## Slide 1
Aqua MODIS Reflective Solar Bands (RSB)
Calibration Improvements in Collection 6
Amit Angal, Junqiang Sun, Hongda Chen, and Jack Xiong
&
MODIS Characterization Support Team (MCST)
Page 1
GRWG/GDWG Annual Meeting 2013-03-04/08

## Slide 2
Outline
Introduction
  Overview
  Reflective Solar Bands (RSB) calibration & Performance
  RSB LUT Timeline
Collection 6 (C6) Algorithm Formulation
  Sensor response (m1)
  Sensor response versus scan-angle (RVS)
C6 Performance
  C5-C6 Comparison
  Improvements to the L1B product (long-term drift, detector-detector striping)
Summary
Page 2

## Slide 3
RSB EV Reflectance
Key Look-Up-Tables (LUT) for RSB
  m1: Inversely proportion to gain at the AOI of SD
  RVS: Sensor Response versus Scan angle (normalized to SD view AOI to the scan mirror)
  Uncertainty
Calibration Sources
  SD/SDSM calibration
  Lunar observation
  SRCA and EV mirror side (MS) ratios
  Selected EV targets
Overview
Angle of Incidence (AOI)
MODIS scan mirror
11.25
Page 3

## Slide 4: MODIS RSB Calibration Using SD/SDSM
Scan Mirror (MODIS)
Sun
1.44% Screen
SDSM
SD
Optional 7.8% Screen
Reflectance Factor
DSD: SD degradation factor;
GSD: SD screen vignetting function
d: Earth-Sun distance
dn*: Corrected digital number; dc: Digital count of SDSM

## Slide 5
MODIS Lunar Calibration

## Slide 6: RSB Performance
Larger SD BRF degradation at shorter (VIS) wavelength
Changes in RSB responses are AOI dependent
Wavelength dependency for gain change
λ
Time
SD
Lunar

## Slide 7
RSB LUT Timeline
m1 LUT History:
  Collection 2: Prelaunch calibration coefficients
  Collection 3-5 : Time-dependent m1 generated using SD/SDSM calibration
  Collection 6:
    m1 generated using SD/SDSM, lunar calibration and EV response trending for Terra bands 1-4,8 and 9 and Aqua bands 8,9
    m1 for all other RSB generated using the SD/SDSM calibration
RVS LUT History:
  Collection 2-3: Prelaunch RVS
  Collection 4: Time-dependent RVS based on SD, lunar and SRCA gains. Applied for bands 3,8 and 9
  Collection 5: Time-dependent RVS applied to all RSB (except SWIR) using SD, lunar and SRCA and EV mirror side differences
  Collection 6: C5 Approach applied for most RSB. EV response trending supplements on-board measurements for Terra bands 1-4,8 and 9 and Aqua bands 8 and 9
Page 7

## Slide 8
Enhancements in Collection 6
Page 8
SD degradation at 936 nm is applied
  Previous SD degradation is normalized at this wavelength
  A correction of 0.6% in Aqua over 11 years (also applied in C5)
Time dependent RVS applied to bands 13-16
  Approach developed to monitor bands 13-16 lunar calibration stability (some pixels saturate when viewing the Moon)
Detector dependent RVS
  Mainly applied to VIS bands (e.g. bands 8-12)
Some RSB calibration coefficients (m1) and RVS are derived at the same time using observations to the SD, Moon, and “pseudo-invariant” EV targets at different AOIs
  Mainly applied to VIS bands (e.g. bands 8-9)
Apply to both Terra and Aqua MODIS

## Slide 9
MODIS RSB RVS characterization
Page 9
Bands with no on-orbit change in RVS (SWIR)
  Pre-launch RVS
Bands with small changes in RVS
  Use SD and lunar trending for mirror side 1 (MS1) RVS
  Use SD, lunar, and EV mirror side ratios for mirror side 2 (MS2) RVS
  Fit each response trending over time, normalize to SD response, and then fit over AOI
Bands with large changes in RVS (Aqua bands 8,9 and Terra bands 1-4, 8 and 9)
  Supplement the on-board lunar measurements with the trending from EV targets
  Detector dependent RVS applied for several VIS bands (B8-12)

## Slide 10: EV Response Trending for VIS bands
A linear approximation is used to estimate the gain change at AOI other than the SD (50.25°) and SV (11.25°)
A long-term bias is also evident for selected VIS bands at the SD AOI
In order to supplement the on-board SD and lunar measurements, EV response trending from pseudo-invariant desert targets was used to track the gain change at all AOI
  Libya-4, Libya-1 and Libya-2

## Slide 11: Results: C5- C6 Difference
Similar improvements observed at other AOI
Long-term bias reduced for all the bands using this approach (Terra bands 1-4, 8, 9 and Aqua bands 8 and 9)
C5 C6
C5 C6
C5 C6

## Slide 12: Results: C5- C6 Difference

## Slide 13: C5-C6 Difference

## Slide 14: Summary
Collection 6 RSB m1 and RVS LUT have been developed for the MODIS L1B products
  Long-term drift observed in Terra bands 1-4, 8-9 and Aqua bands
  8-9 significantly reduced
  Time-dependent RVS applied to bands 13-16
  Detector-dependent RVS for Terra bands 3, 8-12 and Aqua bands 8-12
  C6 RSB uncertainty algorithm provides an improved characterization of the AOI and time-dependence of the RVS
Challenges
  Polarization
  Stability of the EV targets
  Aging instruments: Terra 13+ years and Aqua 11+ years

## Slide 15: Backup Slides

## Slide 16
RSB LUT Summary
Page 16
Terra
Aqua
RVS pl : Pre-launch RVS
RVS* : Time-dependent RVS using SD and Moon for MS#1. MS#2 RVS derived using MS#1 as a reference and using EV and SRCA mirror side ratios.
RVS**: Time-dependent RVS using SD and Moon for both MS#2 (2-point RVS). On-orbit RVS applied since 08/10/2010 in C5 and for entire mission in C6.
RVS***: Time-dependent RVS using SD, SRCA and Moon for MS#1. MS#2 RVS derived using MS#1 as a reference
RVSEV: Time-dependent RVS using the Moon and earth-scene measurements from pseudo-invariant desert targets at various AOI.
Additional correction for the SDSM D9 degradation has been applied to all bands except SWIR starting 01/01/2009.
Detector-dependent RVS applied for Bands 8-12 starting C6.
Version/Bands | Bands 1-4 | Bands 5-7, 26 | Bands 8-12 | Bands 13-16 | Bands 17-19
C4 | m1/RVSpl | m1/RVSpl | m1/RVS*** | m1/RVSpl | m1/RVSpl
C5 | m1/RVS* | m1/RVSpl | m1/RVS* | m1/RVSpl | m1/RVS**
C6 | m1EV/RVSEV | m1/RVSpl | m1EV/RVSEV (B8,9) m1/RVS* (B10-12) | m1/RVS* | m1/RVS**
Version/Bands | Bands 1-4 | Bands 5-7, 26 | Bands 8-12 | Bands 13-16 | Bands 17-19
C4 | m1/RVSpl | m1/RVSpl | m1/RVS** | m1/RVSpl | m1/RVSpl
C5 | m1/RVS** | m1/RVSpl | m1/RVS** | m1/RVSpl | m1/RVS**
C6 | m1*/RVS* | m1/RVSpl | m1EV/RVSEV (B8,9) m1/RVS* (B10-12) | m1/RVS* | m1/RVS**