## Slide 1: GOES Imager/Sounder IR Diurnal Variations
Fangfang Yu and Xiangqian Wu
NOAA/NESDIS/STAR

## Slide 2: Outline
GOES IR midnight Calibration Issue
  Midnight calibration anomaly
  Remedy algorithm – midnight blackbody calibration correction (MBCC) for Imager only
Characterization of the Diurnal Calibration variation
  GEO-LEO
  GEO-GEO

## Slide 3: GOES IR Midnight Calibration Issue
Midnight calibration anomaly
  Last 6-8 hours
  Around the satellite midnight
The radiance in general are well calibrated beyond the midnight effect period using the standard calibration algorithm

## Slide 4: Midnight Calibration Anomaly
Causes of Midnight calibration anomaly (Johnson and Weinreb 1996)
  Extra radiation reflected by the non-unity emissivity of BB to the detectors, when the instrument is viewing the BB.
  Scattered solar radiation contamination at space view
    Most pronounced at pre-eclipse and post-eclipse around the equinox.
Original calibration slopes of GOES-8 Imager Ch2 (from Johnson and Weinreb, 1996)
non-perfect BB effect
Contaminated SPLK effect, mainly at pre- and post- eclipse season around equinox periods
The issue exists at all the current 3-axis stabilized Geo satellites

## Slide 5: Midnight Blackbody Calibration Correction (MBCC)
Current operational remedy: midnight blackbody calibration correction (MBCC) (Johnson and Weinreb1996, Weinreb and Han 2003)
MBCC is based on the regression of the calibration slope against the instrument temperature beyond the midnight effect window
Expected to work better at shortwave channels than at the long-wave channels

## Slide 6: Characterization of Diurnal Calibration Variation
GEO-LEO Inter-calibration
  Reference stable instruments: AIRS and IASI
  Different collocation time period
  Difference between AIRS and IASI radiometric calibration is very small and stable (Wang et al., 2012)
GEO-GEO inter-comparison/inter-calibration
  Frequent collocation over the overlapped area
  Challenge for the SRF correction

## Slide 7: GOES Imager IR channels
Satellite | IR Channel | Central Wavelength (um)
G08-15 | Ch2 | 3.9
G08-11 | Ch3 | 6.75
G12-15 |  | 6.5
G08-15 | Ch4 | 10.7
G08-11 | Ch5 | 12.0
G12-15 | Ch6 | 13.3

## Slide 8: Diurnal Cal. Variation and MBCC Performance Evaluation
Yu et al. 2013, TGRS
Chart: GOES-11 Imager Ch3/4/5 Dirunal Cal. Variation
| Ch3 | Ch4 | Ch5
winter (K) | 0.18 | 0.79 | 0.71
Summer (K) | 0.19 | 0.61 | 0.5
MBCC correction was not always turned on around the midnight time at different channels
assuming

## Slide 9: Diurnal Cal. Variation and MBCC Performance Evaluation
Chart: GOES-12 Imager Ch2/3/6 Dirunal Cal. Variation
| Ch3 | Ch4 | Ch6
winter (K) | 0.54 | 0.66 | 0.73
Summer (K) | 0.61 | 0.64 | 0.6

## Slide 10: MBCC Performance Evaluation – GOES-11 Imager IR Channels
provided by Rama Mundakkara
Ch2 Ch3 Ch4 Ch5
Jan
Feb
Mar
Apr
May
Jun
Jul
Aug
Sep
Oct
Nov
Dec
Jan
Feb
Mar
Apr
May
Jun
Jul
Aug
Sep
Oct
Nov
Dec
Diurnal variation varies at different channel and different season

## Slide 11: GOES Sounder IR instrument characterization
Detector | Channel number | wavelength (µm) | Purpose | Temperature Range (K)
LW (HgCdTe) | 1 | 14.71 | Temperature Sounding | Space - 260
 | 2 | 14.37 | Temperature Sounding | Space - 260
 | 3 | 14.06 | Temperature Sounding | Space - 270
 | 4 | 13.64 | Temperature Sounding | Space - 290
 | 5 | 13.37 | Temperature Sounding | Space - 300
 | 6 | 12.66 | Temperature Sounding | Space - 315
 | 7 | 12.02 | Surface Temperature | Space - 330
MW (HgCdTe) | 8 | 11.03 | Surface Temperature | Space - 335
 | 9 | 9.71 | Total Ozone | Space - 310
 | 10 | 7.43 | Water Vapor | Space - 300
 | 11 | 7.02 | Sounding | Space - 285
 | 12 | 6.51 | Sounding | Space - 265
SW (Insb) | 13 | 4.57 | Temperature Sounding | Space – 310
 | 14 | 4.52 | Temperature Sounding | Space - 295
 | 15 | 4.45 | Temperature Sounding | Space - 275
 | 16 | 4.13 | Temperature Sounding | Space - 330
 | 17 | 3.98 | Surface Temperature | Space - 335
 | 18 | 3.74 | Surface Temperature | Space - 335

## Slide 12
Channel | ΔTb4:00am,0:00am (K) | Channel | ΔTb4:00am,0:00am (K) | Channel | ΔTb4:00am,0:00am (K)
Ch1 | +0.25 | Ch7 | +0.54 | Ch13 | +0.33
Ch2 | +0.22 | Ch8 | +0.49 | Ch14 | +0.38
Ch3 | +0.25 | Ch9 | +0.67 | Ch15 | +0.61
Ch4 | +0.37 | Ch10 | +0.39 | Ch16 | +0.28
Ch5 | +0.24 | Ch11 | +0.44 | Ch17 | +0.51
Ch6 | +0.51 | Ch12 | +0.53 |  | 
Impact of the midnight calibration anomaly for GOES-11 Sounder IR channel 1 through Ch17. ΔTb4:00am,0:00am = ΔTb4:00am - ΔTb0:00am
Yu and Wu, 2013 TGRS
No MBCC applied to the Sounder IR channels

## Slide 13
Channel | max(ΔTb to AIRS) – min(ΔTb to AIRS) (K) | Channel | max(ΔTb to AIRS) – min(ΔTb to AIRS) (K)
Ch1 | +0.10 | Ch7 | +0.51
Ch2 | +0.16 | Ch8 | +0.51
Ch3 | +0.16 | Ch9 | +0.60
Ch4 | +0.22 | Ch10 | +0.36
Ch5 | - | Ch11 | +0.33
Ch6 | +0.46 | Ch12 | +0.35
Impact of the midnight calibration anomaly for theGOES-12 Sounder IR LW and SW (Ch1-12), calculated as the max(ΔTb to AIRS) – min(ΔTb to AIRS) for the night-time collocation data.

## Slide 14
SW Channels | Ch13 | Ch14 | Ch15 | Ch16 | Ch17
Tb difference (K) | +0.60 | +0.63 | +0.91 | +0.60 | +0.76
Tb bias difference between AIRS and IASI around satellite midnight at the SW channels of the GOES-13 Sounder. ΔTb11:00pm,1:00am = ΔTb11:00pm - ΔTb1:00am

## Slide 15
Channel | ΔTb8:00pm,1:00am (K) | Channel | ΔTb8:00am,1:00am (K)
Ch1 | +0.06 | Ch7 | +0.50
Ch2 | +0.13 | Ch8 | +0.55
Ch3 | +0.20 | Ch9 | +0.56
Ch4 | +0.34 | Ch10 | +0.25
Ch5 | +0.36 | Ch11 | +0.23
Ch6 | +0.48 | Ch12 | +0.10
Diurnal calibration variations of GOES-15 Sounder IR channel 1 through Ch12. ΔTb8:00pm,1:00am = ΔTb8:00pm - ΔTb1:00am

## Slide 16: GEO-GEO IR Inter-Calibration
New SRF
1)
2)
- D. Tobin et al. 2005 JGR
This algorithm can be applied to Sounder IR data. However, due to the application of MBCC calibration slope and smoothing of the radiance, it is challenging to apply to Imager IR channels
3)
Where and are the simulated GEO1 and GEO2 radiances using various atmospheric profiles. Could have large uncertainty at the sounding channels

## Slide 17: GEO-GEO Direct Inter-Comparison during GOES-15 PLT Period
GEO-GEO collocation
  Spatial collocation
  Temporal difference: < 5 minutes
  Geometric alignment: viewing zenith angle difference < 1%
GOES 13: -75o
GOES-15: -89.5o

## Slide 18: Direct GEO-GEO IR inter-comparison
BRDF
Midnight
Midnight
Midnight
Midnight
No correction was made for SRF difference
Note: the GEO-GEO inter-comparison was conducted before the SRF shifts for G15 Ch3 and Ch6

## Slide 19: Latitudinal distributions of the mean GEO-GEO Tb difference
Note: the GEO-GEO inter-comparison was conducted before the SRF shifts for G15 Ch3 and Ch6
Significant trending!
No correction was made for SRF difference

## Slide 20: Direct GEO-GEO Inter-Comparison – radiance quality monitoring

## Slide 21: Conclusions
GOES Imager/Sounder IR radiometric calibration is very stable beyond the midnight calibration anomalous period
  Reliable GSICS GEO-LEO IR products to correct the GOES IR radiance.
NOAA applied an empirical calibration algorithm, MBCC method, to correct the erroneous midnight calibration coefficients for the Imager IR channels.
  The MBCC slope was not always triggered on for all the IR channels around the midnight time
  May result in relative larger uncertainty for the midnight IR radiance correction than the day-time data.
Challenges of GEO-GEO product
  Different SRFs
  Bias could be latitudinal dependent.
  The duration of detectable midnight calibration variation may vary at different IR channels. Need to ensure that at least one of the GEO instrument is beyond the midnight impact for an accurate GEO-GEO correction.
Direct GEO-GEO inter-comparison (without SRF correction) is useful for the radiance quality monitoring

## Slide 22
backup

## Slide 23: Scatter-plots of CRTM Simulated G13 vs. G15 Tb values
ECMWF: TB
Courtesy of H. Qian
Notes: Two RTM models with two types of atmospheric profiles used for this study. CRTM model with ECMWF profiles over the collocated region (80W-90W;0-60N), Sampling profiles: 1x1°, 6hr, 3day (Sept. 24-26, 2010), totally around 61x10x4x3=7300 profiles.
LBLRM with 9 profiles.

## Slide 24: Scatter-plots of LBLRM Simulated G13 vs. G15 Tb values
Nine atm. Profiles:
Clear Tropical
Cloud Tropical 700Pa
Cloud Tropical 500Pa
Clear US standard
Cloud Tropical 200Pa
Cloud US standard 500pa
Clear mid. Latitude Summer
Clear mid. Latitude winter
Chart: G13 vs. G15 SRF - Ch6
| 275.63 | 269.70999999999964 | 255.94 | 219.66 | 268.56 | 245.1 | 217.23 | 273.37 | 258.78999999999957
Ch6 | 277.77 | 271.27 | 256.79 | 219.72 | 270.68 | 245.75 | 217.18 | 275.37 | 260.38
Chart: G13 vs. G15 SRF - Ch3
| 248.37 | 248.34 | 246.8 | 219.55 | 241.7 | 238.52 | 217.08 | 247.2 | 242.07
Ch3 | 246.44 | 246.44 | 245.43 | 219.48 | 239.65 | 237.38 | 217.08 | 245.23 | 240.37
Chart: G13 vs. G15 SRF - Ch2
| 300.31 | 281.5 | 261.60000000000002 | 219.73 | 285.95999999999964 | 249.32000000000016 | 216.76 | 292 | 270.45999999999964
Ch2 | 300.05 | 281.32 | 261.48 | 219.92 | 285.71 | 249.23 | 216.98 | 291.76 | 270.24
Chart: G13 vs. G15 SRF - Ch4
| 297.82 | 282.22000000000003 | 262.06 | 219.78 | 286.92999999999961 | 249.75 | 216.73 | 292.01 | 271.56
Ch4 | 297.69 | 282.12 | 261.96 | 219.66 | 286.83 | 249.65 | 216.61 | 291.9 | 271.46

## Slide 25: G13 vs. G15 Imager Ch2 (3.9µm)
Mean Tb at night yet beyond the midnight effect: 0.21K
G13 vs. G15 SRF
Direct inter-comparison
Corrected G13 vs. SRF transformed G15 Tb diff. (CRTM/ECMWF)
Corrected G13 vs. SRF transformed G15 Tb diff. (LBLRM/9 atm.)
Mean Tb at night yet beyond the midnight effect: -0.04 (0.17)K
Yu et al. SPIE 2012

## Slide 26: G13 vs. G15 Imager Ch4 (10.7µm)
G13 vs. G15 SRF
Direct inter-comparison
Corrected G13 vs. SRF transformed G15 Tb diff. (CRTM/ECMWF)
Corrected G13 vs. SRF transformed G15 Tb diff. (LBLRM/9 atm.)
Mean Tb (day-time, 13:00-20:00UTC) : 0.08K
Mean Tb (day-time, 13:00-20:00UTC) : 0.02(0.14)K

## Slide 27: G13 vs. G15 Imager Ch3 (6.5µm)
Mean Tb (day-time) : -0.41K
G13 vs. G15 SRF
Direct inter-comparison
Corrected G13 vs. SRF transformed G15 Tb diff. (CRTM/ECMWF)
Corrected G13 vs. SRF transformed G15 Tb diff. (LBLRM/9 atm.)
Mean Tb (daytime): -0.65(0.06)K
Yu et al. SPIE 2012

## Slide 28: G13 vs. G15 Imager Ch6 (13.3µm)
Mean Tb (day-time, 13:00-20:00UTC) : 0.04K
G13 vs. G15 SRF
Direct inter-comparison
Corrected G13 vs. SRF transformed G15 Tb diff. (CRTM/ECMWF)
Corrected G13 vs. SRF transformed G15 Tb diff. (LBLRM/9 atm.)
Mean Tb (day-time, 13:00-20:00UTC) : 0.32 (0.26)K