## Slide 1
MODIS Response versus Scan-angle (RVS) and Lunar Calibration Reference
Xiaoxiong (Jack) Xiong
NASA/GSFC, Greenbelt, MD 20771, USA
Acknowledgements:
MODIS Characterization Support Team (MCST)
VIIRS Characterization Support Team (VCST)
GSICS Joint Meeting on Research and Data Working Groups, 16-20 March 2015, Delhi, India

## Slide 2
Outline
MODIS Instrument
  Response versus Scan-Angle (RVS)
Solar and Lunar Calibration
On-orbit RVS Characterization
Lunar Calibration Reference
Take-home Message
2

## Slide 3
MODIS: on NASA EOS Terra and Aqua missions launched in 1999 and 2002, respectively
2-sided scan mirror
Spectral range: 36 bands between 0.4 µm and 14.5 µm
  20 RSB and 16 thermal emissive bands (TEB)
Focal plane assemblies (FPA): VIS, NIR, SMIR, and LWIR
Spatial resolution: 250, 500, 1000 m
Swath Width: 2230 km
On-board Calibrators: SD, SDSM, BB, SV, and SRCA
Follow-on instrument: VIIRS on SNPP and future JPSS missions
3
MODIS Instruments

## Slide 4
4
Principal Scan Angles
(Earth View: -55˚ to 55˚)
Angles of Incidence
(Earth View: 10.5˚ to 65.5˚)
NADIR
0˚
Earth Scan
Start (West)
-55˚
Earth Scan
Stop (East)
55˚
SV Port
-98.5˚
On_Board
BB
-129.4˚
SCAN MIRROR
Earth Scan Start 10.5˚
Earth Scan Stop 65.5˚
NADIR 38˚
Normal
SV Port -11.2˚
On_Board BB -26.7˚
MODIS Mirror Scan Angles and AOIs
Pre-launch measurements to characterize sensor response versus scan-angle

## Slide 5
Solar and Lunar Calibration (Aqua MODIS)
SD and SDSM Calibration
  Regularly scheduled with calibration frequency gradually decreased as mission continues to operate beyond its design lifetime (6 years)
  A SD door is close when no calibration is performed (different for Terra MODIS and SNPP VIIRS)
  More than 535 SD/SDSM calibrations performed since launch
Lunar Calibration
  Regularly scheduled via S/C roll maneuvers (8-10 times / year)
  Performed at nearly the same phase angles (-55⁰)
  More than 120 lunar calibrations performed since launch (additional free Moon events)
5

## Slide 6
6
Geometric Factors
gain  1/m1
MODIS SD and lunar observations are made at different AOIs
Solar and Lunar Calibration Methodology
Reflectance factor is computed via a SD calibration coefficient m1

## Slide 7: On-orbit RVS Characterization
Track on-orbit changes in sensor response for each band (l), mirror side (MS) and at different angle of incidence (AOI)
Prior to C6
  SD and lunar trends at 2 different AOIs are used
  Linear interpolation/extrapolation is applied for other AOIs
  EV mirror side ratios are used for some bands (C5) for mirror side 2
Current C6
  Calibration reference traceable to on-board SD
  Relative calibration normalized to lunar calibration
  Changes in sensor response tracked at multiple AOIs, including the use of EV targets (for a few select bands)
7

## Slide 8
Aqua MODIS Radiometric Responses (VIS/NIR)
8
SD Calibration (MS1)
SD Calibration (MS2)
Terra MODIS MS differences are much larger in VIS/NIR

## Slide 9
Aqua MODIS Radiometric Responses (VIS/NIR)
9
SD AOI
Lunar AOI

## Slide 10
Collection-5 MODIS Band-1 EV Reflectance Trends at Various Frames (AOI)
10

## Slide 11
Collection-5 MODIS Band-3 EV Reflectance Trends at Various Frames (AOI)
11

## Slide 12
Collection-5 MODIS Band-9 EV Reflectance Trends at Various Frames (AOI)
12

## Slide 13
Collection-6 MODIS Band-1 EV Reflectance Trends at Various Frames (AOI)
(under science testing)
13

## Slide 14
Collection-6 MODIS Band-3 EV Reflectance Trends at Various Frames (AOI)
(under science testing)
14

## Slide 15
Collection-6 MODIS Band-3 EV Reflectance Trends at Various Frames (AOI)
(already implemented)
15

## Slide 16: Lunar Calibration Reference
16
15
Phase angle dependence of the ROLO model (blue bands)
Comparisons of MODIS and PLEIADES lunar observations with ROLO (unscheduled + scheduled Moon for MODIS)

## Slide 17: Lunar Calibration Reference
17

## Slide 18: Summary
Both Terra and Aqua MODIS instruments continue to operate normally
All on-board calibrators remain capable of their design functions
Dedicated effort for on-orbit RVS characterization is as important as regular SD and lunar calibration (at fixed AOIs)
Overall Aqua MODIS on-orbit performance is better than its predecessor – Terra MODIS
Constant improvements and regular LUT updates are critical to maintaining sensor performance and data quality
Lunar calibration reference (impact due to phase and liberation angle difference)
18