## Slide 1: Development of FY-3/MWRI Calibration on Warm Target and Reflector Emissivity
Shengli Wu@NSMC
XinXin Xie@SSITT
Miao Zhang@NSMC
Hongxin Xie@SSITT
GSICS 2018@Shanghai 2018.03.22

## Slide 2: Outline
Motivation
Warm Calibration system of FY-3/MWRI
Warm Target correction
Reflector correction
Conclusion

## Slide 3: Motivation
Observation minus ECMWF Background Statistics for June 2014
Data: 1 month, after cloud screening, ocean +/- 60 latitude

## Slide 4: Motivation
19V Channel
e.g. channel 3 (19V) O - B
Ascending/Descending Biases:

## Slide 5: Warm Calibration system of FY-3/MWRI
is the efficiency of main reflector;
is the efficiency of hot-reflector
is the efficiency of warm-load
Reflector Emission

## Slide 6: Warm Target correction

## Slide 7: Warm Target correction : Error Source 1/2
Brightness temperature of earth-surface in the direction of the warm reflector back lobe;
Error Source 1
Error Source 2

## Slide 8: Warm Target correction : Error Source 1
Difference of Min/Max BT
Difference of BT result before and after correction
  operational: using a climate brightness Temperature dateset (from AMSR-E Jan-Dec/2006, monthly average);
  Improve: Consider the local time of the warm reflector back lobe direction

## Slide 9: Warm Target correction : Error Source 2
is not the real factor of warm reflector back lobe, the true value should be smaller.
cold
Warm
We can’t use the ground measured to get the back lobe factor directly.

## Slide 10: Warm Target correction : Error Source 2
The parameters that affect scale factors(Gain)
  Temperature of Receiver (Parameter 1);
  Warm/cold reflectors emission(Parameter 2);
  Back lobe of warm reflector(Parameter 3).
Sudden change Only appear on back lobe of warm reflector.
P1
P2
P3

## Slide 11: Warm Target correction : Error Source 2

## Slide 12: Warm Target correction : Error Source 2
Different Color: Back lobe factors
Up: Gain
Down: Stdev of Gain

## Slide 13: Warm Target correction : Error Source 2
Gain
Stdev of Gain
Before
Gain
Stdev of Gain
After

## Slide 14: Warm Target correction : Error Source 2
NeDT(total orbit data)

## Slide 15: Reflector Correction
19V Channel
SSMIS
MWRI

## Slide 16: Warm/cold reflector emission
BT Before and after reflector emissivity correction:
After the reflector emissivity correction, the BT difference is:
]
is the emissivity change
is the cold reflector temperature
is the warm reflector temperature

## Slide 17
降轨
降轨
升轨
升轨

## Slide 18: Change emissivity from 0.9996 to 0.97

## Slide 19: Warm/cold reflector emission
“Real” BT
MWRI BT
Ascending/Descending Bias =warm reflector emission+cold reflector emission
nonlinear
Back lobe

## Slide 20: Warm/cold reflector emission(Emi V1.0)
Y
b
X1
X2

## Slide 21: Warm/cold reflector emission(Emi V 2.0)
Consider the material and technology of warm and cold reflectors, only emissivity of warm reflector could be changed on orbit (or larger than the aluminum emissivity in the very beginning).
For “real” BT, b should be 0
 | Ɛ of pure aluminum | Tested Ɛ (cold sky mirror)
10.65 GHz | 0.00036 | 0.00036
18.7 GHz | 0.00047 | 0.00048
23.8 GHz | 0.00053 | 0.00054
36.5 GHz | 0.00066 | 0.00067
89 GHz | 0.00103 | 0.00104

## Slide 22: Coast Buffer and Scatter filter (Emi V 2.0)
Coast Buffer
Scatter Filter
37V-37H>50 K
37V>19V
19H<185 K
37H<210 K
89V-19V>10 K
89H-19H>30 K
89V>23V
89V>37V
89H-37H>10 K

## Slide 23: 10 Channels’ Regression
Channel | 10.65 | 18.7 | 23.8 | 36.5 | 89
V H | 0.0434 0.0875 | 0.0503 0.0707 | 0.0389 0.0541 | 0.0437 0.0615 | 0.0342 0.0406
Warm reflector emissivity
Ascending/Descending Bias
 | Channel | 10.65 V | 10.65 H | 18.7 V | 18.7 H | 23.8 V | 23.8 H | 36.5 V | 36.5 H | 89 V | 89 H
A/D bias (K) | before Correction | 1.75 | 1.78 | 2.31 | 2.10 | 2.10 | 2.24 | 2.26 | 2.29 | 2.17 | 2.27
 | after correction | -0.004 | -0.008 | -0.008 | -0.007 | -0.008 | -0.006 | -0.004 | -0.001 | -0.008 | -0.014

## Slide 24: Distribution of A/D Bias
FY-3C Bias
FY-3C Emi V 2.0
FY-3C Emi V 1.0

## Slide 25: Conclusion
After the Warm Target correction, stability of the Gains of MWRI’s 10 channels are significant improved.
Reflector correction reduced the ascending/descending Bias from 1~2K to less than 0.1K.

## Slide 26: Thank You for your attention!