## Slide 1: 空白演示
HIRAS status update and IR recalibration for FY-3/IRAS and VIRR
Chengli Qi, Chunqiang Wu, Hanlie Xu, Xiuqing Hu, Na Xu, Lu Li
National Satellite Meteorological Center(NSMC), CMA
Mingjian Gu, Dekun Yin
Shanghai Institude of Technical Physics(SITP), CAS
GSICS IR Group web meeting, 19 Mar 2020

## Slide 2
Outline
  FY-3D/HIRAS status updates
  Recalibration for FY-3/IRAS
  Recalibration for FY-3/VIRR
  Future work plan

## Slide 3
1. FY-3D/HIRAS status updates -1
Instrument heating and decontamination
1st-Dec, 2018
2nd-Oct, 2019
Heating operation on Oct 21, 2019
LW and MW bands signal are nearly recovered.
SW band is slight weaker in signal
ISRF compared with best signal status

## Slide 4
SNO BT comparison with JPSS-1/CrIS
general less than 0.3, 0.7 and 1.0 K in the LWIR, MWIR and SWIR bands, respectively.
Consistent with the accuracy of before heating.

## Slide 5
Cold space view angle adjustment
1. FY-3D/HIRAS status updates -2
Carminati. F, Xiao. X, Lu.Q, et al. Assessment of the Hyperspectral Infrared Atmospheric Sounder (HIRAS)，Remote Sens, 2019, 11, 2950, doi: 10.3390/rs11242950
Assessment from Met Office
abnormal negative deviations of Detector 3 in the middle wave band in some areas

## Slide 6
Raw DS spectra of 4 Fovs in SW band
monthly analysisof the deep space and blackbody observation spectra
deep space spectra of FOV3 in the middle wave (MW) and short wave (SW) bands were abnormal: they rose sharply during January to March 2019

## Slide 7
20190201 SW波段4个FOV观测DS原始光谱与纬度和太阳高度角之间的关系
Correspondence between anomalous spectra and latitude, longitude and solar height angle with typical channels : LWIR: 778cm-1; MWIR: 1300cm-1; SWIR: 2108cm-1.
SW Detector 3 is abnormally high in deep space observation in the range of 90-120 ° of the solar zenith angle, which leads to the abnormal deviation in calibration bias.

## Slide 8
Cold Space observation spectra : 10 view angles, -94.93°,-91.30°,-87.69°,-84.09°,-80.48°,-76.89°,-73.29°,-69.70°,-66.10°,-62.50°
nominal CS angle at -71°
DS spectra at angles of -94.93 °, -91.30 °, -87.69 °, -84.09 °were not affected by solar contamination
Deep space pointing angle was decided to set from -71 ° to -87 °
After adjustment
Before adjustment
FOV2
FOV1
FOV3
FOV4

## Slide 9
distribution of RTTOV simulations and observed brightness temperature deviations (O-B) before and after the adjustment of deep space scan angle, respectively
O-B deviation of the solar contamination area deep space is consistent with that in other areas after the DS observation angle adjustment, the O-B bias abnormal isbasically eliminated,
Global Distribution of Deviations between Observed and RTTOV Simulation (FOV3) Before and after Deep space scan Angle Adjustment

## Slide 10
IRAS instrument:
Infrared Atmospheric Sounder
Onboard FY-3A/B/C
20 thermal infrared channels of observation in range of 3.7~15 um , 6 NIR and VIS channels between 0.69~1.64 um.
Atmospheric temperature and water vapor retrieval .
NWP & Climate research.
2. Recalibration for FY-3/IRAS
Recalibration program object:
Build 10-year infrared sounder recalibration datasets (IRAS).
FY-3A 0.5 year, FY-3B 6 year, FY-3C 8 year
3 versions, V1, V2 and V3.
V1: system bias corrected. V2: calibrated data using fine recalibration model. V3: FCDR dataset, platform difference removed and grided products.

## Slide 11
Calibration accuracy evaluation method and data
SNO method: FY-3B/C cross with Metop-A/IASI
SNO Pairing Method
Time Constrains:
10 minutes
Air Path Constrains
FOV 26-31 for iras
Distance constrains
5 km
Scene uniformity
IRAS: std/btm < 0.001
dataset
FY3B/IRAS 2012-2018 7 years
FY3C/IRAS 2014-2018 5 years

## Slide 12
FY-3B/IRAS time series bias characteristics
Ch2 14.71um
Ch3 14.49um
Ch12 7.33um
Ch1 14.95um
Ch16 4.47um
Ch17 4.45um
Most of channels bias within 0.5 K.
Absorbing channels show seasonal variation. Different characteristics in day and night time samples.
Day time samples
Night time samples

## Slide 13
FY-3C/IRAS time series bias characteristics
Ch2 14.71um
Ch3 14.49um
Ch11 7.43um
Ch1 14.95um
Ch16 4.47um
Ch17 4.45um

## Slide 14
Day time samples
Night time samples
Operational Version
Test V1
Test version: system bias corrected based on IASI measurement, abnormal bias corrected due to code bugs.
Bias std are almost same from operational to test version datasets.

## Slide 15
Day time samples
Night time samples
Operational Version
Test V1
Test version: Absorbing channels large system bias slightly corrected .
Bias std are almost same from operational to test version datasets.

## Slide 16
Using Metop-A/IASI as reference, the historical bias of TE channels (@ 250K) are assessment based on SNO inter-calibration method.
10.8 um
12 um
VIRR-IASI
Systematic bias are showed on all VIRRs；
Inconsistency between different series;
Inconsistency between two channels on the same satellite ；
Strong seasonal variation;
Especially on FY-3A
3. Recalibration for FY-3/VIRR Thermal channels

## Slide 17: Scene dependence of FY-3/VIRR
The bias's scene dependence of FY-3A/B (Jul. 2012) and FY-3C/VIRR (Dec. 2015)
Strong nonlinearly in FY-3A/VIRR:
10.8 um: on low temperature, VIRR is much cooler than IASI;
12um: on low temperature, VIRR is much warmer than IASI;
In-flight calibration method: Linear calibration and nonlinear correction
Nonlinear correction coefficients maybe changed in orbit, so the nonlinearity of TE bands is not described exactly, leading to bias’s scene dependence and seasonal variation;
FY-3A/VIRR: nonlinear correction based on GSICS
FY-3BC/VIRR: systematic correction based on GSICS
FY-3A
FY-3B
FY-3C

## Slide 18: Nonlinear coefficient correction of FY-3A/VIRR based on GSICS
Ref.
Radiation
FY linear Radiation
Sample selection
Quadratic regression
New nonlinear coefficients
New nonlinear correction
New FY Radiation
(Xu et al., 2014)
(Xu et al., 2014)
Using new nonlinear coefficients on other sample data:
Before
After
Before
After
10.8um
12 um

## Slide 19: Correction result of FY-3A/VIRR
Before
After
10.8um
12 um

## Slide 20: Systematic correction results of FY-3B/VIRR
Seasonal variation and nonlinearity are not strong at FY-3B;
Using SNO collocation samples to do systematic correction;
10.8um
12 um
Samples: using 2012 and 2013 SNO samples based on GSICS, Metop-A/IASI as reference;
Regression: Calculating corrected coefficients using linear regression;
After correction, the consistency between two TE channels are much more better;

## Slide 21: Conclusion and future work
FY-3D/HIRAS is in stable operational status.
HIRAS MW and SW bands FOV3 solar contamination was eliminated by adjustment of cold space viewing angle from -71° to -87 °.
Long-term accuracy evaluation results are stable for much of channels of IRAS, except for strong absorbing ch1,13,15,16,17,18 and they present different seasonal variations for day and night time samples.
V1 IRAS and VIRR recalibration datasets generated and validated.
Future work will focus on the development of refined physical infrared recalibration model and V2 datasets.