## Slide 1: 2020 Annual GSICS Calibration Report - GEO
Chengli Qi
Na Xu, Xiuqing Hu, Hanlie Xu,
Qiang Guo, Boyang Chen, Lin Sun
National Satellite Meteorological Center(NSMC), CMA
2020 GSICS Webmeeting,
16, Apr, 2020

## Slide 2: 2019 Geo Calibration work in CMA
Major satellite/Instrument events of FY-2 & FY-4 relevant to calibration and the impact on the radiance quality
FY-4/GIIRS status updates in 2019 and performance of V3 L1 data.
Radiometric calibration performance for FY-2 & FY-4 Imagers using GSICS inter-calibration/inter-comparison
  Table/Figure summarizing bias/unc: average of last one year or typical time period
  Time series chart of bias/unc: since operation start or last one year
  Scene Dependent Bias: Scatterplot of scene dependent bias

## Slide 3: FY-2 Calibration Events in 2019
http://planet.rssi.ru
FengYun satellite operation information and announcements:
http://www.nsmc.org.cn/en/NSMC/Channels/100029.html
Satellite (status) |  | Current Location | Launch Date | Operation Date
FY-2E | (11, Jan 2019 disengaged) |  | 23, Dec 2008 | Feb 2009
FY-2F | （Op） | 112°E | 13, Jan 2012 | 
FY-2G | （Op） | 105°E | 31, Dec 2014 | 1, Jul 2015
FY-2H | （Op） | 79°E | 5, Jun 2018 | 11, Jan 2019
Major Calibration Related Events | FY-2F | FY-2G | FY-2H
New calibration algorithm (source free) implementation | January, 2020 | January, 2019 | 

## Slide 4: FY-4 Calibration Events in 2019
http://planet.rssi.ru
Satellite (status) |  | Current Location | Launch Date | Operation Date
FY-4A | (Op) | E104.7 | 2016-12-11 | 
Major Calibration Related Events | FY-4A/AGRI
Improvement of De-stripping algorithm | 08/11/2019
reflective solar bands (RSB) calibration correction coefficients updates | 09/25/2019

## Slide 5
FY-2G/H IR Rad. Calibration Performance (2019)
 | IR1_11.0 (286K) | IR2_12.2 (285K) | IR3_6.9 (247K) | IR4_3.8 (285K)
Metop-A/IASI | -0.2593 | -0.7841 | -1.4897 | 1.1197
Metop-B/IASI | -0.2194 | -0.7426 | -1.4335 | 0.9641
 | IR1_11.0 (286K) | IR2_12.2 (285K) | IR3_6.9 (247K) | IR4_3.8 (285K)
Metop-A/IASI | -0.4519 | -1.0307 | -1.1022 | 2.9536
Metop-B/IASI | -0.4653 | -1.0320 | -0.9186 | 1.8630
FY-2G
FY-2H
IR1
IR2
IR3
IR4
IR1
IR2
IR3
IR4
Ref: IASI-A
Ref: IASI-A
Bias at window channels are less than 1k (FY-2G) and 1.5k (FY-2H)
FY-2G/H’s Tb bias at standard Tb show seasonal variation;

## Slide 6
FY-4A/AGRI IR Rad. Calibration Performance
 | CH8_3.8 (286K) | CH9_6.25 (237K) | CH10_7.1 (248.5K) | CH11_8.5 (283K) | CH12_10.8 (286) | CH13_12.0 (285) | CH14_13.5 (258)
Metop-A/IASI | -3.5927 | -0.2230 | 0.6527 | 0.3048 | -0.1227 | -0.1447 | -0.2635
Metop-B/IASI | -3.5148 | -0.2223 | 0.6303 | 0.2745 | -0.1419 | -0.1776 | -0.2756
*Average bias in 2019;
CH 9
CH 10
CH 11
CH 12
CH 13
CH 14
Tb bias are less than 0.7 k (operation phase);
Scene dependent is much more better than FY-2G/H;
No strong seasonal variation;

## Slide 7: Improvement of radiance consistency in different detectors
Wave Length (um) | Sensitivity | Stripe Noise
3.5L | 0.1K | 
5.8 | 0.09K | Visiable in image
6.9 | 0.12K | 
8.0 | 0.06K | Visialbe in products
10.3 | 0.07K | 
11.5 | 0.14K | 
13.2 | 0.6K | Visiable in image
Stripping in absorption channels are more obvious, main contribution from inconsistent SRFs.
The maximum equivalent noise can reach 2K, improved algorithm of adaptive wavelet filter has been preliminarily developed.
The requirement of quantification is getting higher, systematic design needs to be strengthened and the image quality needs to be improved continuously
13.2-13.8um

## Slide 8: Improvement of radiance consistency in different detectors
SST
Cloud Mask
Before de-stripping
After de-stripping

## Slide 9
FY-4A AGRI RSB
Monitoring and calibration updates using stable target
band | Wavelength (um) | Total attenuation rate(%) | Annual attenuation rate(%)
1 | 0.47 | 40.3 | 14.36
2 | 0.65 | 6.62 | 2.36
3 | 0.82 | -8.56 | -3.05
5 | 1.61 | 6.53 | 2.33
6 | 2.85 | 9.62 | 3.43
Operational Ref bias variation in 2 years;
Normalized response trend using Multi method;
Reflection of Ch1 at DCC and Multi station are consistent;

## Slide 10: DCC monitored reflectance of operational VS recalibrated
DCC reflectance (operational)
DCC reflectance (recalibrated)
15% fluctuation in Operational calibration results
Fluctuation was greatly decreased in recalibrated results

## Slide 11: Milestones of FY-4/GIIRS L1 Data (Operation, Distribution and Application)
2016.12~2017.10: Commissioning and Preliminary Validation
2017.11~2018.04: User Readiness (i.e. observation mode optimization, training)
2018.05~2018.11: Experimental Application (i.e. NWP and retrieval)
2018.12: Operational Application in GRAPES (direct radiation assimilation)
2019.01: V1 Data Distribution to International Users;
2019.08: V2 Data in Operation(QC，ZPD alignment)
2019.10: V2+ available (spectral coding error corrected)
2019.11.07: V3 Data in Operation (diurnal BT biases variation corrected)

## Slide 12
Key Improvements of GIIIRS L1 Data in V3
Abnormal Observations (bad graduals) Correction
Spectral Optimization
  Off-axis effects correction (for 32×4 detector array)
  Accurate spectral shift correction
Radiometric Optimization
  ZPD-aligning improvement
  New calibration model for diurnal-varying on GEO platform
Processing Efficiency Optimization
  Fast DFT for GIIRS utilization (increased by 8~10 times)
Optional Apodization upon L1 Data
  0813~0828: Hamming Apodization
  0829~ : Unapodized

## Slide 13
Bad granules occurred in V1 has been overcome completely: 11/350≈3%
V2/V3 V1

## Slide 14: FY4A/GIIRS spectral and radiometric asessment
ECMWF Assessments of V3 Radiance (O-B): channel7 (~100hPa)
V1
V3
Less than 1K for most channels in LW band
SSEC/UW Assessments of V3 Spectral Shifts
< 5ppm
< 10ppm

## Slide 15: Future work
Add GSICS monitoring of FY-4/GIIRS using GSICS inter-calibration/inter-comparison
  Bias at reference temperature
  Uncertainty of Tb bias at reference temperature
FY-4A/AGRI IR gsics correction products development
FY-4A/AGRI reflective solar bands (RSB) calibration monitoring.