## Slide 1: GEO-GEO analysis of Himawari-8/9 parallel observation
OKUYAMA Arata and KODERA Kazuki
Meteorological satellite center of Japan meteorological agency

## Slide 2: Introduction
Himawari-9
Himawari-8 is currently a back up satellite (no observation operation)
Specifications of H8 and 9 are same.
  140.7 deg. East
  Imager is AHI.
  16 bands, FD obs. every 10 mins. and area obs. every 2.5 mins.
Launch | 02 Nov. 2016
Back up operation start | 10 Mar. 2017
Parallel data dissemination start | 27 Sep. 2022
Full operation start | 13 Dec. 2022
Himawari-8/AHI
Himawari-9/AHI
01
03
02
04
06
05
07
11-08
16
15
14
13
12

## Slide 3: AHI9 calibration (IR bands, B07-B16)
+0.2
+0.1
0.0
-0.3
-0.2
-0.1
-0.4
The brightness temperature (Tb) biases for both AHI8 and AHI9 are stable and within 0.3K at the standard radiance corresponding to clear sky sea surface temperature.
There are some cases that have larger biases for lower Tb.
AHI9 | AIRS | N20/CrIS | IASI-B | IASI-C
B07 (3.9) | 0.09 | - | -0.12 | -0.11
B08 (6.2) | - | -0.17 | -0.20 | -0.20
B09 (6.9) | -0.09 | 0.03 | 0.00 | 0.00
B10 (7.3) | -0.20 | -0.12 | -0.15 | -0.14
B11 (8.6) | - | - | -0.09 | -0.08
B12 (9.6) | -0.20 | -0.12 | -0.21 | -0.17
B13 (10.4) | -0.04 | -0.02 | -0.06 | -0.05
B14 (11.2) | -0.05 | -0.03 | -0.07 | -0.06
B15 (12.4) | -0.04 | -0.06 | -0.10 | -0.09
B16 (13.3) | -0.10 | -0.15 | -0.18 | -0.17
The bias wrt. CrIS at B07, the bias wrt. AIRS at B08 and the bias wrt. AIRS and CrIS at B11 are omitted because of its large uncertainty.
2022/09/02 - 30
AHI8 | AIRS | N20/CrIS | IASI-B | IASI-C
B07 (3.9) | 0.04 | - | -0.13 | -0.13
B08 (6.2) | - | -0.16 | -0.19 | -0.19
B09 (6.9) | -0.33 | -0.21 | -0.24 | -0.23
B10 (7.3) | -0.22 | -0.12 | -0.14 | -0.14
B11 (8.6) | - | - | -0.07 | -0.06
B12 (9.6) | -0.25 | -0.19 | -0.28 | -0.24
B13 (10.4) | 0.06 | 0.07 | 0.03 | 0.04
B14 (11.2) | 0.07 | 0.08 | 0.04 | 0.05
B15 (12.4) | 0.01 | -0.02 | -0.05 | -0.04
B16 (13.3) | 0.15 | 0.10 | 0.07 | 0.09
AHI - reference [K]
Tb bias at the standard radiance [K]

## Slide 4: Tb bias stability at the std.rad.
B07
B08
B09
B10
B11
B12
B13
B14
B15
B16
cf. GSICS Calibration Products Plotting Tool
2022/09
2022/12
2022/09
2022/12
AHI9 Tb bias at std. rad.
10 weeks from 2022/09
wrt. Metop-B/IASI (Re-analysis)
AHI8 Tb bias at std. rad.
10 weeks from 2022/09
wrt. Metop-B/IASI (Re-analysis)
The biases for AHI9 are as stable as AHI8, and their range is similar to the one for AHI8.
+0.1
-0.4
+0.1
-0.4

## Slide 5: Tb bias stability at 220K
2022/09
cf. GSICS Calibration Products Plotting Tool
AHI9 Tb bias at 220K
10 weeks from 2022/09
wrt. Metop-B/IASI (Re-analysis)
2022/12
2022/09
2022/12
AHI8 Tb bias at 220K
10 weeks from 2022/09
wrt. Metop-B/IASI (Re-analysis)
B07
B08
B09
B10
B11
B12
B13
B14
B15
B16
+0.4
-1.5
Their variation ranges are almost same for the both of AHI8 and 9, while there are some cases that have larger value.
+0.4
-1.5

## Slide 6: Tb bias (Diurnal variation ) at the std. rad.
AHI8
1.0
-1.0
AHI9
B07
B08
B09
B10
B15
B16
B13
B14
B11
B12
B07
B08
B09
B10
B15
B16
B13
B14
B11
B12
There is no significant variation for the both imagers.
x-axes: Time in UTC
y-axes: Tb bias [K]
Note that the bias wrt. CrIS at B07, the bias wrt. AIRS at B08 and the bias wrt. AIRS and CrIS at B11 have large uncertainty due to their band gaps.
Data period:
2022/09/02 - 30
1.0
-1.0
1.0
-1.0
1.0
-1.0
1.0
-1.0
AIRS
CrIS-20
IASI-B
IASI-C

## Slide 7: Tb bias (for scan position by GSICS approach)
Grid ave. Tb difference residual: GEO minus LEO [K]
B07
B09
B08
B10
B11
+0.5 K
-0.5 K
2022/10 to 12
3 month average
AHI-9 vs. LEO
Only scenes with Tbs > Std.radiances are averaged.
Hyper-sounder Tb
AHI9 - Hyper-sounder Tb
B12
B14
B13
B15
B16
B12
B14
B13
B15
B16
vs.
Metop-B/IASI
vs.
SNPP/CrIS
Scan position dependent bias is not significant.

## Slide 8: GEO-GEO: Himawari-8/9
Some bands show difference brought by the SRF difference between H8 and H9.
B07: Solar light reflection affects stronger for B07 (3.9um) of AHI9 than the one of AHI8, due to a difference of SRFs.
B16: Their SRFs doesn't look different comparing with B07's case, however, observed Tb is affected by the SRF difference, e.g. the standard radiances are 269.73K and 268.53K for AHI8 and AHI9, respectively.
Note that the GEO-GEO differences don't mean absolute error but just a relative comparison between AHI9 and AHI8.
2022/09/01 00:00 UTC
B07
B09
B08
B10
B11
B12
B16
B15
B14
B13
+5 K
Tb difference: snapshot (AHI9 - AHI8) [K]
-5 K
B07
B16
AHI8
AHI9
AHI8
AHI9
SRF difference is not corrected. Images are compared by "as is".

## Slide 9: GEO-GEO (Trend)
Tb difference: AHI9 - AHI8 [K]
Trend of the full disk average of the Tb difference is stable in the last a few month, while there is diurnal variation.
B07: diurnal variation due to the contribution by the solar light component.
B16: larger difference with around 1.5 K due to the SRF difference.
B09: relatively larger difference than other bands. GSICS approach revealed that Himwari-8 has around 0.3K of bias on this band.
In other bands, the difference between the both imagers are within less than 0.1K.
B07
B16
B09
2022/10
2022/12
2022/11
2022/10
2022/12
2022/11

## Slide 10: GEO-GEO (Diurnal variation)
Amplitude of the diurnal variation wrt. AHI8 is around order of 10-2 K except for B07 and B16.
AHI8 and AHI9 change a position of space-look for calibration reference before 03UTC and 15UTC.
0.5
-0.01
-0.05
0.37
0.35
0.02
0.01
-0.01
-0.02
-0.04
-0.06
-0.06
-0.07
-0.03
-0.04
0.00
-1.40
-1.50
Tb difference: AHI9 - AHI8 [K]: 00 to 23 UTC
B07 (3.9)
B09 (6.9)
B08 (6.2)
B10 (7.3)
B11 (8.6)
B12 (9.6)
B16 (13.3)
B15 (12.4)
B14 (11.2)
B13 (10.4)
2.0
0.02

## Slide 11: GEO-GEO (Diurnal variation)
B13 (10.4um)
AHI9 – AHI8 [K]
2022-09-12, 00 to 11 UTC
Local noon is around 03 UTC
SRF difference is not corrected.
Both satellites are located at “140.7 degrees East”.
-0.2
0.0
03 UTC
02 UTC
00 UTC
01 UTC
07 UTC
06 UTC
04 UTC
05 UTC
11 UTC
10 UTC
08 UTC
09 UTC

## Slide 12: GEO-GEO (Diurnal variation)
0.0
+0.2
03 UTC
02 UTC
00 UTC
01 UTC
07 UTC
06 UTC
04 UTC
05 UTC
11 UTC
10 UTC
08 UTC
09 UTC
B15 (12.4um)
AHI9 – AHI8 [K]
2022-09-12, 00 to 11 UTC
Local noon is around 03 UTC
SRF difference is not corrected.
Both satellites are located at “140.7 degrees East”.

## Slide 13: GEO-GEO (Diurnal variation)
Larger Tb bias (positive or negative) around 0.2K are seen over land area in daytime afternoon in “window” bands (B13, 14, 15, and slightly B11).
According to GEO-GEO liner regression coefficients (AHI-9 vs. AHI-8), the Tb difference corresponding to observed Tb is estimated as the bottom figure.
The estimation can roughly express the Tb change, but further quantitative check will be necessary.
Estimated Tb diff. (AHI9-AHI8) [K]
Observed Tb [K]
B13
B14
B15

## Slide 14: GEO-GEO (for scan position)
Faint systematic difference can be seen in East-West direction for B08, B11, B12 and in center-to-limb direction for B16.
Grid ave. Tb difference residual: AHI9 - AHI8 [K]
+1 K
-1 K
+1 K
-1 K
+0.1 K
-0.1 K
+0.1 K
-0.1 K
+0.1 K
-0.1 K
+0.1 K
-0.1 K
+0.1 K
-0.1 K
+0.1 K
-0.1 K
+0.1 K
-0.1 K
+0.1 K
-0.1 K
2022/11/01-10
10 days average
B07
B09
B08
B10
B11
B12
B16
B15
B14
B13
Hyper-sounder Tb
AHI9 - Hyper-sounder Tb

## Slide 15: summary
Himawari-9 started full operational observation on 13 December 2022 at 140.7 deg. East. AHI9 specification is the same as that of AHI8.
Himawari-8 has been a back up satellite since 13 December 2022.
The Tb biases for both of AHI8 and AHI9 are within 0.3K at the standard radiance.
The Tb bias trend and diurnal variation for AHI9 are stable in similar order to the one of AHI8, and no significant scan position dependency.
GEO-GEO comparison with AHI8 shows diurnal variation, trend, and scan position dependency. B07 and B16 show relatively larger variation. Potential cause is the difference of SRFs.
Again, the “GEO-GEO” differences don't mean absolute bias but just a relative comparison between AHI9 and AHI8.

## Slide 16: References
The Level-1 products comparison between H8 and 9
  https://www.data.jma.go.jp/mscweb/en/oper/healthcheck/[L1]Validation_H9_processing_forMSCWEB.pdf
Calibration portal
  https://www.data.jma.go.jp/mscweb/en/oper/calibration/calibration_portal.html
Event log
  https://www.data.jma.go.jp/mscweb/en/oper/event_H9.html

## Slide 17
Back up…

## Slide 18: Tb bias (for scan position by GSICS approach)
Grid ave. Tb difference residual: GEO minus LEO [K]
B07
B09
B08
B10
B11
+0.5 K
-0.5 K
2022/10 to 12
3 month average
AHI-9 vs. LEO
Only scenes with Tbs > Std.radiances are averaged.
Hyper-sounder Tb
AHI9 - Hyper-sounder Tb
B12
B14
B13
B15
B16
B12
B14
B13
B15
B16
vs.
Metop-B/IASI
vs.
SNPP/CrIS
Scan position dependent bias is not significant.
AHI-8

## Slide 19
B13 (10.4um)
AHI9 [K]
(not a difference)
2022-09-12
00 to 11 UTC
Local noon is around 03 UTC
Tb over Australia varies from around 290 to 320K.

## Slide 20: Linear regression (sorry, night time only)
AHI8
AHI9
B07
B09
B08
Brightness temperature [K]
B11
B10
B13
B12
B16
B15
B14
Stable linear relationship since the start of AHI-9 observation in 2022
Data period:
1st - 7th October 2022
Conditions:
10 – 13 UTC (Nighttime)
Sun zenith angle: larger than 110 deg.
Data from outside the Earth disk not included
https://www.data.jma.go.jp/mscweb/en/oper/healthcheck/[L1]Validation_H9_processing_forMSCWEB.pdf

## Slide 21: Linear regression (sorry, night time only)
 | Slope | Offset
B01 | 1.04750 | -0.19260
B02 | 1.00594 | -0.03052
B03 | 0.97746 | -0.02018
B04 | 0.99698 | -0.00105
B05 | 0.94478 | -0.01444
B06 | 0.98394 | 0.00502
 | Slope | Offset | H9-H8 [K] @StdTb.* | H9-H8 [K] @220K | H9-H8 [K] @250K | H9-H8 [K] @290K
B07 | 0.99118 | 2.70964 | 0.19 @285.95K | 0.77 | 0.51 | 0.15
B08 | 0.99759 | 0.55558 | -0.01 @234.65K | 0.03 | -0.05 | -0.14
B09 | 1.00486 | -0.80853 | 0.38 @243.85K | 0.26 | 0.41 | 0.60
B10 | 0.99805 | 0.51075 | 0.02 @254.59K | 0.08 | 0.02 | -0.05
B11 | 0.99893 | 0.28697 | -0.02 @283.82K | 0.05 | 0.02 | -0.02
B12 | 0.99738 | 0.63064 | -0.05 @259.45K | 0.06 | -0.02 | -0.13
B13 | 1.00129 | -0.41459 | -0.05 @286.18K | -0.13 | -0.09 | -0.04
B14 | 0.99766 | 0.6151 | -0.05 @286.10K | 0.10 | 0.03 | -0.06
B15 | 1.00284 | -0.75661 | 0.05 @283.78K | -0.13 | -0.05 | 0.07
B16 | 0.97016 | 6.29606 | -1.75 @269.73K | -0.27 | -1.17 | -2.36
Std.Tb: Brightness temperature corresponding to clear-sky sea surface (band-dependent)
Data period: 1st – 7th October 2022, 00-06UTC (for B01-B06), 10-13UTC (forB07-B16)
Linear regression coefficients
https://www.data.jma.go.jp/mscweb/en/oper/healthcheck/[L1]Validation_H9_processing_forMSCWEB.pdf