## Slide 1: Applications of GIRO Models on the Calibration of GOES Imager Visible Channel
Fangfang Yu and Xiangqian Wu
12/02/2014-12/04/2014
Notes: Changed GIRLO to GIRO

## Slide 2: GOES Imager Visible Channel
GOES-E
GOES-W
Channel name | Central wvlen | Detector number | IGFOV | IGFOV at nadir (earth) | IGFOV at Moon Surface | E-W oversampling rate | Downlink resolution
1 | 0.65um | 8 | 28µrad | 1km | ~11km | 1.75 | 10-bits
GOES15(West): 105W
GOES12(East): 75W

## Slide 3: Moon Acquisition
Unscheduled Moon
  Full Disk scan
  Northern Hemisphere Extended (GOES-East)
  Northern Hemisphere (GOES-West)
  Only a few good data per year
Scheduled Moon
  Since Oct. 2005
  Using the 1-minute of star obs. window
  About one or two times per month
Courtesy of : http://www.ospo.noaa.gov/Operations/GOES/east/routine_strategy.html
Notes: Removed “after”

## Slide 4: Moon Images
Unscheduled Moon image:
  ~2-6 good moons per year
Scheduled Moon Image
  1-2 times per months, if not in conflict with Hurricane tracing event
  P1 scan using the star obs. time slots
Space Calibration
Earth

## Slide 5: Phase Angle Range
Phase angle ranges
  unscheduled moon: Could be any angle between [-90o, 90o], appear in the space area within GOES FOR.
  Scheduled moon: [-75o, 75o] except for [-7o, 7o], 0.5o away from the Earth edge

## Slide 6: GOES Imager Visible Calibration
No on-orbit calibration device
On-orbit calibrated with cold reference (space)
  Clamp at space at every 2.2 sec (FD) or 36.6 sec (other mode)
  Reset the zero-radiance output to a predetermined value of 29 raw count

## Slide 7: GOES Lunar Irradiance
Where n is the number of pixels associated with the moon illumination, ωi is the solid angle subtended by moon pixel i, fi is its oversampling factor, Ci is the pixel (i) output signal in count, Ci,bk is the background space signal in count, and Si is the prelaunch calibration coefficient to convert instrument signal output from raw counts to radiance. Since the pre-launch calibration coefficients for the eight detectors of GOES imagers visible channel are similar, their medium value is used to calculate the lunar irradiance (http://www.star.nesdis.noaa.gov/smcd/spb/fwu/homepage/GOES_Imager_Vis_PreCal.php).
ω=7.18x10-10Sr-1
S = constant
f = 1.75

## Slide 8: Background Count – 1) Stray-light
OOF response near the Moon
OOF response near the Earth

## Slide 9: Background counts – 2) High Energy Particles
“Pepper & Salt” pixels
  Cosmic rays, rare cases
Other noises
  Very rare cases

## Slide 10: Background Count – 3) Calculation
150pixels
10 pixels of interval
Mean count value of the un-contaminated space pixels in between the ellipse belt
150 and 160 pixels from the moon edge
~17,000pixels

## Slide 11: Oversampling
Constant value of 1.75 (ITT et al. 2010 GOES-NOP Data Book)
  IFOV = 28urad
  Scan rate (20 degree/second) x detector sample rate (21840 samples/second) = 16urad
  E-W Oversampling factor = 28/16 = 1.75
  N-S: oversampling factor = 1

## Slide 12: Integration
i: pixels within the 150-pixel ellipse area
S: pre-launch calibration coefficient of the representative detector

## Slide 13: Special handling….
Unclipped moon data
  Too close to the Earth edge (>0.5 degree from the Earth edge)
  Challenging to derive the fitting ellipse automatically
Phase angle = 0.73o/1.02o
0.22o
0.47o

## Slide 14: GIRO Results
Fitting function: y = a + bt + ct^2
Trending agrees very well with those of DCC and Ray-match methods
  Deviations of Desert method results in 2004-2005 and 2009-2010 are apparent
Larger uncertainty than DCC and ray-match methods

## Slide 15: Comparison with Integrated Method Results
Phase angle dependent calibration uncertainty ?
?
?
Yu, F. and X. Wu, 2014, RSE, in review

## Slide 16: Discussions
Major Irr. Measurement Uncertainty Components
  Oversampling factor: constant = 1.75?
    INR(IMC uncertainty)
    Moon-satellite relative movements
      E-W
      N-S
    Instrument performance: space clamp
  Background count
    Cosmic rays and other noises
    Stray-light from the Earth
    Detector noise
  Moon irradiance measurement
    Stray-light from the illuminated moon pixels
    MTF
    Detector bright target recovery
    Detector noise

## Slide 17: Oversampling factor
where Dmoon-sat is the distance between the Moon and GOES satellite, ω is the solid angle extended from the detector (7.84x10-10Sr-1), m is the number of illuminated moon pixels, rmoon is the moon radius (1735.5km), and fractionmoon is the fraction of illuminated moon, and ϕ is the selenocentric elongation of the Earth from the Sun.
Key number!

## Slide 18: Illuminated Moon Boundary
How to determine the illuminated moon boundary?
Sub-pixel illumination
Stray-light
MTF effect
Detector bright target recovery effect
Method 1: boundary identification

## Slide 19: Illuminated Moon Pixels
Method 2 - Histogram
How to determine the threshold between stray-light contaminated space pixels and dark moon pixels, especially for the large phase angle images?
20pixels from the fitting ellipse

## Slide 20: Future Instrument Cal. Spec.
GOES-R ABI
  6 VNIR bands
Radiometric Calibration Accuracy Request
  3% absolute calibration accuracy @ 100% albedo
  0.5% stability over the mission life
Moon phase angle range at [-90o, 90o]
Courtesy of X. Shao and C. Cao
Band | Central wavlen(µm)
1 | 0.47
2 | 0.64
3 | 0.86
4 | 1.38
5 | 1.61
6 | 2.25

## Slide 21: Needs for Lunar Model Development
Irradiance model (ROLO/GIRO)
  Spectral irradiance
    And how to derive channel irradiance from spectral irradiances, current method and potential improvement.
  Quantify absolute calibration accuracy at difference phase angles
    To meet the 3% abs. calibration accuracy specification
    Plan for improvement
  Quantify relative calibration accuracy at different phase angle
    To meet the 0.5% stability specification