## Slide 1: Beyond the workshop: Developing inter-calibration products based on Lunar Observations
Tim Hewison (EUMETSAT) [GSICS Research Working Group Chair]
Special Issue of the IEEE TGRS on “Inter-Calibration of Satellite Instruments”:

## Slide 2: Overview
Introduction to GSICS
Scoping Lunar Inter-Calibration products
Choice of Reference Instrument
Error Budgets
Key Issues
  Phase Angle-Dependence
  Spectral Consistency
Moving Forward

## Slide 3
Global Space-based Inter-Calibration System
What is GSICS?
  Global Space-based Inter-Calibration System
  Initiative of CGMS and WMO
  Effort to produce consistent, well-calibrated data from the international constellation of Earth Observing satellites
What are the basic strategies of GSICS?
  Improve on-orbit calibration by developing an integrated inter-comparison system
        Initially for GEO-LEO Inter-satellite calibration
        Being extended to LEO-LEO
        Using external references as necessary
  Best practices for calibration & characterisation
This will allow us to:
  Improve consistency between instruments
  Reduce bias in Level 1 and 2 products
  Provide traceability of measurements
  Retrospectively re-calibrate archive data
  Better specify future instruments
EUMETSAT | CNES | JMA
NOAA | CMA | KMA
ISRO | NASA | WMO
USGS | NIST | JAXA
ROSHYDROMET | IMD | ESA
Notes: 3
Well, what is GSICS?
The Global Space-based Inter-Calibration System (GSICS) is an initiative of CGMS and WMO, which aims to ensure consistent calibration and inter-calibration of operational meteorological satellite instruments.
One of the basic strategies of GSICS is to develop an integrated on-orbit cal/val system - initially by LEO-GEO Inter-satellite/inter-sensor calibration.
This will then allow us to provide corrections to improve the consistency between instruments, produce less biased level 1, and therefore, level 2 data, and ultimately, retrospectively re-calibrate archive data, which is of great interest in the climate monitoring community

## Slide 4: GSICS Organisation
GSICS Exec Panel
GSICS Coordination Center
GSICS Research Working Group
GSICS Data Working Group
VIS/NIR Sub-Group
Microwave Sub-Group
UV Sub-Group
Future Sub-Groups...
WGCV IVOS
WGCV MWSG
GPM X-CAL
WGCV ACSG
CEOS ACC
CGMS
WMO
Identified potential collaborations with CEOS

## Slide 5
GSICS Principles
TRACEABILITY / UNBROKEN CHAINS OF COMPARISONS
Systematic generation of inter-calibration products
  for Level 1 data from satellite sensors
  to compare, monitor and correct the calibration of monitored instruments to community references
  by generating calibration corrections on a routine operational basis
  with specified uncertainties
  through well-documented, peer-reviewed procedures
  based on various techniques to ensure consistent and robust results
Delivery to users
  Free and open access
  Adopting community standards
To promote
  Greater understanding of instruments’ absolute calibration, by analysing the root causes of biases
  More accurate and more globally consistent retrieved L2 products
  Inter-operability for more accurate environmental, climate and weather forecasting products

## Slide 6
GSICS Bias Monitoring
  Routine comparisons of satellite radiances against reference
GSICS Correction
  Function to correct issued radiances for consistent calibration with reference
GSICS Calibration
  Alternative Calibration Coefficients = operational calibration + inter-calibration corrections
GSICS Reports & Guidelines
  Recommendations to modify processing/practices
  Design and Operation of future satellite instruments
For Operational Environmental Satellites
   Infra-red recalibration (GEO and LEO)  Pre-Operational & Demo status
  (current operational satellites)  Near real-time and re-analysis
   Visible and near-infrared recalibration (GEO and LEO)  In development within GSICS
   Microwave – Conical & Cross-track Scanners (LEO)  In development with GPM XCAL
   Historic Instruments  In development for SCOPE-CM
GSICS Product Types

## Slide 7: Scoping GSICS Lunar Inter-Calibration Products
Double differences of irradiances from monitored and reference instruments, w.r.t. a model: (IMON -Imod)-(IREF -Imod)
  Assumption: Model errors will cancel out (linearity)
  No need to adjust to absolute scale
In general: Correction=Polynomial function of Radiance, ΔL=g(LMON)
  Simplest case: Linear scaling of operational gain
Allow GSICS Corrections to vary with time
  To track calibration drifts of monitored instrument
  Defined from combination of data over extended period with negligible drift

## Slide 8: Future Extensions to GSICS Lunar Inter-Calibration Products
Combine with other methods
  Aim to provide users ONE product
  For all VIS/NIR/SWIR channels
  Cover full dynamic range (linearity)
  e.g. DCC + Lunar
  Optimal combination weighted by uncertainty of each method
Spectral consistency
  Inter-channel Bias Matrix
Tie to absolute standards
  SI Traceability
Lunar Disc Reflectance Range

## Slide 9: Choice of Inter-Calibration Reference
Common primary reference for all GSICS Products in Reflected Solar Band
Initially, Aqua/MODIS
  For all channels?
Migration to Other Reference Instruments
  e.g. SNPP/VIIRS
  Handled using concept of delta Corrections
  Developed for GEO-LEO IR
  Based on double-difference of different inter-calibration products
  Ensures traceability back to primary reference
  Ensures radiometric consistency
  Eventually hyperspectral reference instruments? TBC!
Needs robust and consistent uncertainty estimates
  Both on observed and modelled irradiances

## Slide 10: Error Budgets
Needed for comparison and optimal combination of
  Observed Irradiances
  Modelled Irradiances
Construct Error Budget (Type B evaluations) for uncertainties due to
  Systematic Biases
  Random Errors
Based on sensitivity analysis of
  GIRO inputs
  Observed Irradiance Calculation
Validate Error Budgets (Type A evaluations)
  For Systematic Parts – by comparison between different instruments
  For Random Parts - By repeated evaluations over range of conditions (and time!)
Need to repeat analysis for new ROLO versions

## Slide 11: Previously on this Channel...
From 2014 GRWG Annual Meeting:
  Sophie Lacherade (CNES) “Moon synergy" 3mb
    Uncertainties for Moon cross-calibration due to the lack of correct spectral interpolation between PHR and MODIS (cross-calibration limited to a band-to-band cross-calibration)
    Estimation of the interband accuracy based on Moon acquisitions ≈ 3%
  Tom Stone (USGS) “Using the Moon as Cross-Calibration reference” 3n
Xiong et al., “Comparison of MODIS and PLEIADES Lunar observations” Proc. SPIE 9241
    calibration differences between MODIS and PLEIADES are general less than 2% with an exception of about 3% for the NIR channel

## Slide 12: Key Issues
Solar Variability
  Impact on time series comparisons
Phase Angle Dependence
  Quantify uncertainty
  Independently for each wavelength
  Distinction from non-linearity
  Distinction from stray light
Spectral Consistency
  Investigate Different possible spectral interpolation methods
  Including Apollo sample smoothing
  Quantify uncertainty
  Inter-channel Bias estimation
Implement improvements of ROLO model

## Slide 13: Moving Forward
Confirm what are the Key Issues?
  Here by Thursday
Who will investigate each of the Key Issues?
  Identify on Thursday
Who will generate Error Budgets?
  Identify on Thursday

## Slide 14: Revised Key Issues
Oversampling factor
Phase Angle Dependence
  Quantify uncertainty
  Independently for each wavelength
  Distinction from non-linearity
  Distinction from stray light & Space Count Subtraction
Spectral Consistency
  Investigate Different possible spectral interpolation methods
  Including Apollo sample smoothing
  Quantify uncertainty
  Inter-channel Bias estimation
  Use of hyperspectral Moon observations to validate SBAFs (Hyperion?)
Implement improvements of ROLO model

## Slide 15
Discussions during the workshop
Over-sampling: what do we mean by over-sampling? How to estimate it?
Cross-comparison of instruments with similar SRFs + similar illuminations conditions when viewing the Moon
Can we put the lunar irradiance on a absolute scale?
Can we think of inter-calibration and tying target instruments to a common reference? If yes, how?
    How do we address the spectral differences between instruments?
    How do we address differences in the illumination conditions between instruments?

## Slide 16: Use of hyperspectral Moon observations to validate SBAFs
Calibration with ROLO accounts for instruments’ SRFs
Inter-calibration using broad-band reference instrument (e.g. MODIS)
Requires Spectral Band Adjustment Factor
  To convert radiances from Reference to be consistent with Monitored Instrument
  In general SBAF are scene-dependent
    Implies inter-calibration would be application-dependent
  Moon irradiance spectrum relatively smooth
  Are Lunar SBAFs applicable to other scenes? E.g. Similar to those for DCCs?
Courtesy of Lawrence Ong

## Slide 17: Moving Forward
Who will investigate each of the Key Issues?
Who will generate Error Budgets?