## Slide 1: Diurnal Cycles and Midnight Calibration Correction
NOAA GPRC
Fangfang Yu, Xiangqian Wu and Haifeng Qian

## Slide 2: Midnight calibration anomaly and correction
Causes of Midnight calibration anomaly (Johnson et al. 1996)
  Extra radiation reflected by the non-unity emissivity of BB to the detectors, when the instrument is viewing the BB.
  Scattered solar radiation contamination at space view
    Most pronounced at pre-eclipse and post-eclipse around the equinox.
Current operational remedy: midnight blackbody calibration correction (MBCC) (Johnson et al. 1996, Weinreb and Han 2003)
Figure. Original calibration slopes of GOES-8 Imager Ch2 (from Johnson et al. 1996)
non-perfect BB effect
Contaminated SPLK effect

## Slide 3: Diurnal Cal. Variation and MBCC Calibration Evaluation
(Yu et al. 2011)
Time dependent Tb bias and the standard deviation (at 30 minute time bin), as well the onset frequency of MBCC application for the winter of 2008.
Chart: GOES-11 Imager Ch3/4/5 Dirunal Cal. Variation
| Ch3 | Ch4 | Ch5
winter (K) | 0.18 | 0.79 | 0.71
Summer (K) | 0.19 | 0.61 | 0.5
Chart: GOES-12 Imager Ch2/3/6 Dirunal Cal. Variation
| Ch3 | Ch4 | Ch6
winter (K) | 0.54 | 0.66 | 0.73
Summer (K) | 0.61 | 0.64 | 0.6

## Slide 4: MBCC correction residual
Channel name | GOES11 - AIRS |  | GOES12 -AIRS | 
 | Winter | Summer | Winter | Summer
Ch3 | 0.12 | 0.13 | - | -
Ch4 | - | 0.49 | - | 0.48
Ch5 | 0.13 | - |  | 
Ch6 |  |  | 0.04 | -0.11
Impact of one of MBCC slopes on the diurnal calibration variations, calculated using Equation 9. (mean Tb 1:30pm – Tb 1:30am). The results with high frequency of MBCC onset (>80%) are in bold-italic font.
Efficiently and sufficiently applied at Ch2
When sufficiently applied, it can work efficiently at Ch3, Ch5 and Ch6 to reduce the Tb bias to less than 0.15K
Even sufficiently applied, it works less efficiently at Ch4 with residual (mean Tb difference to the day-time data) of 0.5K bias.
(Yu et al. 2011)

## Slide 5: Time-dependent correction during the midnight effect time
Calibration uncertainty
  Duration and amplitude vary at different IR channel
  Depends on the frequency of MBCC application
  Change with seasons
  Has seasonal long-term trending
Evaluation of MBCC performance for each channel and in each month for GOES-11
Time-series of Tb bias to AIRS at midnight time for GOES-11 Ch4

## Slide 6
Diurnal Cycle – MET-9
No gap-filling applied

## Slide 7: Spatial Distr. Of Coll. Scenes
Local Satellite Time: 5:00-7:00
Day-time
Need more investigation for the root cause of ~6:00 anomaly

## Slide 8: MTSAT-2

## Slide 9
No gap filling method applied
Different SRF applied for AIRS/IASI spectral transform
Gap filling for the missing AIRS spectra
Difference in the AIRS and IASI calibration
Uncertainty in collocation(?)

## Slide 10
MTS02 : Jan

## Slide 11
MTS02 : Feb

## Slide 12
MTS02 : Mar

## Slide 13
MTS02 : Apr

## Slide 14
MTS02 : May

## Slide 15
MTS02 : Jun

## Slide 16
MTS02 : Jul

## Slide 17
MTS02 : Aug
Ch1

## Slide 18
MTS02 : Sep

## Slide 19
MTS02 : Oct

## Slide 20
MTS02 : Dec

## Slide 21: Time-dependent GSICS Correction
If the diurnal variation is caused by the instrument calibration, then
  Fundamental correction with physical models
  Empirical correction
    Look-up table at time bin
      Large uncertainty for the extreme Tb scenes
    Regression correction at time bin
      LEO overpass time interval (1 hour) vs. BB calibration time interval (30 minutes).
      Need sufficient of collocation pixel number for the all-sky correction at each time bin
Delta correction
  Use SNO and double-difference to transfer the AIRS correction to IASI with uncertainty analysis.

## Slide 22
GOES-11 Ch4
18:00
18:30
19:00
19:30
20:00
20:30
21:00
21:30
22:00
22:30
23:00
23:30
00:00
00:30
01:00
01:30
02:00
02:30
03:00
03:30
04:00
04:30
05:00