## Slide 1: Feasibility of using Deep Convective Clouds to Monitor Performance of the Geostationary Lightning Mapper (GLM)
Dennis Buechler

## Slide 2: Monitoring GLM Performance
GLM has no onboard calibration
Use Lightning Imaging Sensor (LIS) observations as proxy data base
  Similar design
  14 years of data
  Lightning detection efficiency depends on background radiance
Background radiance calibration
  Monitor background calibration stability

## Slide 3
LIS Instrument design
Fast lens, narrowband filter at 777.4 nm
128 x 128 CCD array, 500 fps imaging
Frame-to-frame subtraction isolates lightning transients against bright daytime background
3.3 km nadir to 10 km at corners
Backgrounds ~35 seconds

## Slide 4: Deep Convective Cloud Technique (DCCT) Applied to LIS Background Radiances
Purpose: Determine if Deep Convective Cloud Technique (DCCT) can be used as a stable target for monitoring stability of GLM to meet Measurement Precision Requirements

## Slide 5: Methodology
For each July and August from 1998 through 2010:
Convert LIS observed counts for each background pixel into radiance values
Get adjusted LIS radiance (ρ)
  adjust to SZA of 0° (divide by cosine of SZA)
  assume Lambertian surface
Determine coincident LIS and VIRS pixels

## Slide 6: Methodology (Cont’d)
Use pixels that meet following criteria:
  Tb < 205K
  SZA < 40°
  VZA < 40°
  10° < RAA < 170°
  σ/ρi < .02
  No lightning nearby
Assume constant solar radiance throughout the period
Assume random distribution of VZA, RAA, and SZA

## Slide 7: TRMM LIS observations of Deep Convective Clouds
Super Typhoon Haitang
VIRS IR
LIS BG

## Slide 8: LIS BG DCC Radiance Distribution for July and August (1998-2010)

## Slide 9: Trend of LIS BG DCC radiances

## Slide 10: Land/Ocean LIS BG DCC Distribution

## Slide 11: DCCT Conclusions
The DCCT method to monitor radiance measurement precision is successfully applied to TRMM LIS and VIRS for the period 1998-2010
  The maximum yearly deviation is ~0.6 %
  The maximum deviation occurs in 2000
  Trend is not evident
  Instrument appears stable and well within 5% over its lifetime
The DCCT method appears applicable to monitor GOES-R GLM radiance measurement precision using ABI and GLM
Land pixels appear slightly brighter than ocean pixels

## Slide 12: Additional
Develop and apply an angular distribution model
Examine other stable targets for calibration (e.g., deserts, glint)