## Slide 1: Diurnal Calibration Variation of GOES Infrared Channels
Fangfang Yu and Xiangqian Wu

## Slide 2: Outline
Issue
  Midnight calibration anomaly
  Empirical correction algorithm – midnight blackbody calibration correction (MBCC)
MPCC performance: GOES diurnal calibration variation
Time-dependent calibration correction using GSICS

## Slide 3: GOES IR Calibration Issue
Midnight calibration anomaly
  Expected before GOES-8 was launched
  First observed at short-wave channel with the SST temperature at the Gulf of Mexico
  Impact on the long-wave channel product was also report
  Last 6-8 hours, around the satellite midnight

## Slide 4: Midnight calibration anomaly and correction
Causes of Midnight calibration anomaly (Johnson and Weinreb, 1996)
  Extra radiation reflected by the non-unity emissivity of BB to the detectors, when the instrument is viewing the BB.
  Scattered solar radiation contamination at space view
    Most pronounced at pre-eclipse and post-eclipse around the equinox.
Empirical correction algorithm: midnight blackbody calibration correction (MBCC) (Johnson et al. 1996, Weinreb and Han 2003)
Figure. Original calibration slopes of GOES-8 Imager Ch2 (from Johnson et al. 1996)
non-perfect BB effect
Contaminated SPLK effect

## Slide 5: GOES-11 Imager IR Diurnal Cal. Variation
Time dependent Tb bias and the standard deviation (at 30 minute time bin), as well as the frequency of MBCC application in the winter of 2008 (Jan. 1 – March 1).
Chart: GOES-11 Imager Ch3/4/5 Dirunal Cal. Variation
| Ch3 | Ch4 | Ch5
winter (K) | 0.18 | 0.79 | 0.71
Summer (K) | 0.19 | 0.61 | 0.5
(Yu et al. ITGRS, 2012, revised)

## Slide 6: GOES-12 Imager IR Diurnal Cal. Variation
Chart: GOES-12 Imager Ch2/3/6 Dirunal Cal. Variation
| Ch3 | Ch4 | Ch6
winter (K) | 0.54 | 0.66 | 0.73
Summer (K) | 0.61 | 0.64 | 0.6
(Yu et al. ITGRS 2012, revised)
Time dependent Tb bias and the standard deviation (at 30 minute time bin), as well as the frequency of MBCC application in the winter of 2008 (Jan. 1 – March 1).

## Slide 7: Tb bias to AIRS at satellite midnight
Duration and amplitude vary at different IR channels
Depends on the frequency of MBCC application, which may change with time
May have long-term variation at some IR channels
Evaluation of MBCC performance for each channel and in each month for GOES-11
Time-series of Tb bias to AIRS at midnight time for GOES-11 Ch4

## Slide 8: Possible Solutions…
Fundamental correction with physical models
  Many benefits with this approach, but may need a relatively long term effort
Time-dependent correction using GSICS

## Slide 9: Time-dependent correction using GSICS
Look-up table at each time bin
  Large uncertainty for the extreme Tb scenes
Regression correction at time bin
  Bias is scene radiance dependent, need sufficient collocated scenes
Delta correction
  Difference between the reference instruments, assuming the difference is constant
    SNO
    Double difference using satellite observation or RTM simulation as the transfer
Time interval
  LEO overpass time interval (1 hour) vs. BB calibration time interval (30 minutes).
  Varying interval

## Slide 10
GOES-11 Ch4
18:00
18:30
19:00
19:30
20:00
20:30
21:00
21:30
22:00
22:30
23:00
23:30
00:00
00:30
01:00
01:30
02:00
02:30
03:00
03:30
04:00
04:30
05:00

## Slide 11: How to handling ….
Insufficient collocated scenes within one-month smoothing window (used for the day-time data correction)
  Keep the smoothing period, relax the time interval
  Keep the time interval, extend the smoothing period

## Slide 12
Overlaping period:
  use one reference instrument
  take advantage of more collocated scenes