## Slide 1: GOES-East DCC analysis
Fangfang Yu, Xiangqian Wu and Haifeng Qian
GSICS Annual Meeting in Beijing, 2012

## Slide 2
GOES Imager 4KM data
Subset GOES data
Potential DCC identification
Lat = [-25., 25.]
Lon = [nadir_lon – 30, nadir_lon + 30.]
Time duration = [1:30pm LST – 2.0, 1:30pm LST + 2.0]
SZN < 50
VZN < 50
Tb4 < 215K
STDV(Tb4) < 10K
STDV(Refl) < 5%
Data Analysis
Time titude
Longitude
Solar zenith angle
Solar azimuth angle
Satellite viewing zenith angle
Satellite viewing azimuth angle
Raw count for the solar reflected channel(s)
Radiance for the IR channels
Standard deviation of the reflectance (3x3 pixels window)
Standard deviation of the Tb (3x3 pixels window)
Archived variables in the intermediate products

## Slide 3
Sensitivity Analysis
Monthly DCC pixel
Gain
SCIAMACHY
GOME
BRDF, SRF dependent?
Spectral Conversion
ADM Correction
PDF
Mode/Median
Binsize?

## Slide 4: GOES Imager SRF
GOME DCC Reflectance
MODIS SRF

## Slide 5
SCIAMACY = 0.998, STDV=0.25%
GOME = 0.994, STDV=0.27%
Diff =0.4%
SBAF

## Slide 6: Reflectance Difference from Mcidas Raw vs. Albedo Outputs
Both reflectances are normalized with solar zenith angle and sun-earth distance
Possible cause of the difference might come from the mcidas conversion from raw count to pre-launch radiance/reflectance
Message home: when using the mcidas-x function to read the SRF-dependent variables (e.g., albedo and Tb ), be sure that the Mcidas version installed at the computers are most recently updated!
Notes: Message home: be aware of the mcidas-x Version! Updated SRFs, pre-launch cal. Coefficients?

## Slide 7: DCC PDF ADM vs. NO ADM

## Slide 8: Bin size, Median vs. Mode?
Baseline algorithm:
SZN < 40
VZN < 40
Latitude = [-20, + 20.]
Longitude = [nadir_lon – 20., nadir_lon + 20.]
Tb4 < 205K + 1.15K
STDV(Tb4) < 1K
STDV(Refl) < 3%
Binsize = 1.0%
Binsize = 0.5%
Binsize = 0.2%

## Slide 9: Sensitivity to STDV(Tb10.7um) – Dec11
Binsize=1.0, stdv(refl)<3%
Tb4 < 205K

## Slide 10: Sensitivity to STDV(Tb10.7um) – Jan12
Binsize=1.0, stdv(refl)<3%
Tb(10.7um)<205 + 1.15K

## Slide 11: Sensitivity to STDV(Tb4) - Feb12
Binsize=1.0, stdv(refl)<3%
Tb(10.7um)<205+1.15K
Notes: DCC reflectance (Mode) looks stable with stdv(Tb4)>1.0

## Slide 12: Sensitivity of GOES GVAR Radiance Count
Change in Tb / GVAR Count

## Slide 13: Sensitivity to Tb(10.7um) threshold – Dec11
Binsize=1.0, stdv(refl)<3%
Stdv(Tb10.7um)<1K

## Slide 14: Sensitivity to Tb(10.7um) threshold – Jan12
Binsize=1.0, stdv(refl)<3%
Stdv(Tb10.7um)<1K
Notes: Binsize can affect the fitting function uncertainty

## Slide 15: Sensitivity to Tb4 threshold – Feb12
Binsize=1.0, stdv(refl)<3%
Stdv(Tb10.7um)<1K
Notes: Median refl seems provide more accurate relative calibration accuracy

## Slide 16: Conclusions
Progressive decreasing of DCC reflectance can be observed with mode reflectance (binsize=1.0%) and median reflectance
  Bin size may affect the uncertainty of the gain trending
  median reflectance is irrelevant to the bin-size and more consistent at stdv(Tb10.7um)
While mode/median reflectance might be sensitive to the Tb(10.7um) thresholds, mode has larger stdv and median is more sensitive to the threshold.
When using mode reflectance, we need to consider the impact of instrument noise at cold scenes
  For the 10-bit GOES instrument, it is recommended to slightly relax the criteria of Stdv(Tb10.7) (eg. =1.5K)

## Slide 17: Preliminary Results – Comparison of the Vicarious Calibration Methods
GOES-10 Imager Visible Channel

## Slide 18: Backup Slides

## Slide 19
Feb
Mar
Apr
May
Jun
Oct
Dec
Sep
Nov
MTSAT-2