## Slide 1: Migrating from GSICS GEO-LEO IR Products from Metop-A/IASI to Metop-B/IASI Reference
The advantages of using multiple references
The problems of using multiple references
The solutions to using multiple references!

## Slide 2: Why Migrate References?
Nothing lasts forever (except GSICS!)
  Satellites (and their instruments) have finite life spans
  Satellite programs have longer life spans, but still finite
“There can be only one!”
Better, faster, cheaper* instruments come along
  (*Choose 2 of the above)

## Slide 3: Combining Multiple References
Advantages
Robustness
  In case of failure of one reference
  Allows transition between references – e.g. MetopA->B
Greater coverage of diurnal cycle
  Both scene and instrument calibration variability
  Important for 3-axis stabilized spacecraft
Broader engagement of GSICS members
Disadvantages
Segal's law:
  "A man with a watch knows what time it is. A man with two watches is never sure." [Bloch, Arthur (2003). Murphy's Law. Perigee. p. 36.]
  It refers to the potential pitfalls of having too much conflicting information when making a decision.
  User confusion
Not Metrologically Traceable
  as the references can never be perfectly consistent
Define only one as the calibration reference
All others are regarded as calibration transfer standards

## Slide 4: GSICS Correction using Reference 1
GSICS Correction for Monitored Instrument
Based on Reference 1
Function with uncertainty

## Slide 5: GSICS Correction using Reference 2
GSICS Correction for Monitored Instrument
Based on Reference 2
Different function with different uncertainty

## Slide 6: Delta Correction Reference 2->1
GSICS Correction for Reference 2
Based on Reference 1
Function with uncertainty

## Slide 7: Delta Correction Reference 2->1
Delta GSICS Correction for Reference 2
Based on Reference 1
Function with uncertainty

## Slide 8: GSICS Correction using Reference 2->1
Add Delta Correction to Reference 2
Maps GSICS Correction based on Reference 2 to Reference 1
Consistent function, but larger uncertainty

## Slide 9: Implementation
GSICS Correction
Correction using Reference 1
Correction using Reference 2
Delta Correction Reference 2-1

## Slide 10: GEO as Transfer Radiometer
Direct Comparison may not be possible
GEOs can be collocated with every LEO overpass
Use as transfer standard
But define Delta Corrections in GEO channel space
Errors from Spectral Band Adjustment tend to cancel out
Wang et al [20011] compared direct SNO method with LEO-GEO-LEO method:
  Broad consistency
  Mean differences of <0.1K
  L. Wang, M. Goldberg, X. Wu, C. Cao, R. A. Iacovazzi Jr, F. Yu, and Y. Li, "Consistency assessment of Atmospheric Infrared Sounder and Infrared Atmospheric Sounding Interferometer radiances: Double differences versus simultaneous nadir overpasses," JGR, vol. 116, no. 11, 2011.
Schematic diagram showing how double differencing against a third sensor as an intermediate transfer reference can be used to inter-calibrate two instruments without requiring direct comparison of their observations. Dashed red lines show collocated observations from pairs of instruments. Black arrows show calibration transfers.

## Slide 11: Meteosat-9/SEVIRI as Transfer
Direct Comparison Metop-B & -A impossible
  Orbits are 50min/180° out of phase
Meteosat-9/SEVIRI can be collocated with overpasses of both Metop’s
  Use as transfer standard
Define Delta Corrections in SEVIRI channel space
  Strictly Meteosat-9/SEVIRI
  But SRF Differences expected to have negligible impact (TBC)
Schematic diagram showing how double differencing against a third sensor as an intermediate transfer reference can be used to inter-calibrate two instruments without requiring direct comparison of their observations. Dashed red lines show collocated observations from pairs of instruments. Black arrows show calibration transfers.
Metop-A/IASI
Metop-B/IASI
Meteosat-9/SEVIRI
(Met9/SEVIRI-MetopB/IASI)-(Met9/SEVIRI-MetopA/IASI)
(Met9/SEVIRI-MetopB/IASI)
(Met9/SEVIRI-MetopA/IASI)

## Slide 12: Time Series of Standard Biases (Met9/SEVIRI-MetopB/IASI) & (Met9/SEVIRI-MetopA/IASI)
Change in IASIb processor

## Slide 13: Statistics of Double Difference (Met9/SEVIRI-MetopB/IASI)-(Met9/SEVIRI-MetopA/IASI)
Meteosat-9/ SEVIRI Channel | Mean Difference of Standard Biases [K] | Std Dev of Difference of Standard Biases [K]
IR3.9 | +0.006 | 0.022
IR6.2 | +0.028 | 0.009
IR7.3 | +0.025 | 0.029
IR8.7 | +0.014 | 0.021
IR9.7 | -0.036 | 0.019
IR10.8 | -0.025 | 0.022
IR12.0 | -0.017 | 0.020
IR13.4 | -0.000 | 0.010
Biases for standard scene radiances (clear sky)
From Demonstration GSICS Re-Analysis Corrections
From first 2 months data from Metop-B/IASI
Double differences are small (<|40mK|) for all channels
Only statistically significant at 95% level for IR6.2
  which was unusually stable during this short 21d period
Good News!
  But still need delta correction!
  With uncertainties!
Total Uncertainty (k=1) from Uncertainty Analysis ~0.01K

## Slide 14: Statistics of Double Difference (Met9/SEVIRI-MetopB/IASI)-(Met9/SEVIRI-MetopA/IASI)
Meteosat-9/ SEVIRI Channel | Mean Difference of Standard Biases [K] | Std Dev of Difference of Standard Biases [K]
IR3.9 | -0.72 | 0.28 (<<u)
IR6.2 | +0.11 | 0.05
IR7.3 | +0.15 | 0.18
IR8.7 | -0.11 | 0.30
IR9.7 | -0.17 | 0.08
IR10.8 | -0.30 | 0.29
IR12.0 | -0.25 | 0.24
IR13.4 | +0.67 | 0.28
Biases for cold cloud radiances (220K)
From Demonstration GSICS Re-Analysis Corrections
From first 2 months data from Metop-B/IASI
Double differences are small (<|1K|) for all channels
Statistically significant only for:
  IR6.2 which was unusually stable during this short 21d period
  IR3.9, which is very non-linear at the cold end
More Good News!
Total Uncertainty (k=1) from Uncertainty Analysis ~0.7K for IR3.9 ~0.2K for IR10.8

## Slide 15: Formulating the Delta Correction
GSICS Correction
LMON=g(LREF)=a+b.LREF
Met9-IASIa: LMet9=a1+b1.LIASIa
Met9-IASIb: LMet9=a2+b2.LIASIb
=> Double Difference:
a2+b2.LIASIb=a1+b1.LIASIa
LIASIb=(a1-a2) /b2+(b1/b2).LIASIa
LIASIb=a21+b21.LIASIa

## Slide 16: Propagation of Variances
? terms are unknown
  Could assume them to be zero
  Implies no correlation between MSG2-MetopA/IASI and MSG2-MetopB/IASI
Let’s try…

## Slide 17: Time Series of Double Differences (Met9/SEVIRI-MetopB/IASI)-(Met9/SEVIRI-MetopA/IASI)

## Slide 18: Propagation of Variances
Assuming off-diagonal terms to be zero
  Implies no correlation between MSG2-MetopA/IASI and MSG2-MetopB/IASI
Uncertainty calculated theoretically from the propagation of variance tends to underestimate the observed variability of the time series of RACs
  for both standard scene radiances
  as well as cold scenes (Tb=220K).
This implies the correlation is significant
  and cannot be neglected
  Although it is not trivial to calculate,
  Because two time series are independent
  not necessarily on the same dates
Correlation between RAC coefficients?
  Generally low
  although not independent…