## Slide 1: Desert Calibration for GOES Imager Visible Channels
Fangfang Yu and Xiangqian Wu
NOAA/NESDIS/STAR
2013 GSICS Annual Meeting
Williamsburg, VA
March 4-8, 2013

## Slide 2: Motivations
Long-term stable desert can be used to monitor sensor degradation trending and facilitate cross-comparisons
If well characterized, can serve as a SI-traceable reference.
The only desert that can be observed by both GOES-West (135W) and GOES-East (75W)
Historical reference target for GOES-8/9/10 operational calibration
  Operational target switched to MODIS since 2004 due to the change of desert reflectance (ENSO effect)
  Recent studies indicate the desert is stable at decadal scale
Revisit the Sonoran desert calibration for the GSICS project
  Traceable to Aqua MODIS C6
  Combined method

## Slide 3: Challenges for the long-term Instrument Inter-Calibration
Target should be long-term stable, well-characterized, traceable to community best reference (Aqua MODIS C6 data) with low uncertainty
  Spatially
  Spectrally – smooth target reduces the SBAF uncertainty
  Temporal – various scales, from diurnal to long-term
  BRDF
  Ground target – atmospheric component impacts
Trending within the mission
  Degradation pattern(function)
  BRDF
Transfer the trending to absolute calibration (reference data)
Bridging the gap between different sensor
  Spectral band adjustment factor (SBAF)

## Slide 4: GOES-8 through G15 Imager Ch1 SRFs

## Slide 5: Sonoran Desert
[32.05N-32.35N, 114.4W-114.7W]
25 pixels
Study area
49 pixels

## Slide 6: Satellite Data
Satellite | Satellite Location | Observation time (local) | Data Period
GOES-8 | GOES-East, 75W | ~12:35pm | 11/01/1994 – 04/01/2003
GOES-9 | GOES-West, 135W | ~12:15pm | 09/30/1995 – 07/07/1998
GOES-10 | GOES-West, 135W | ~12:15pm | 07/07/1998 – 06/21/2006
GOES-11 | GOES-West, 135W | ~12:15pm | 06/21/2006 - 12/06/2011
GOES-12 | GOES-East, 75W | ~12:35pm | 04/01/2003 – 04/14/2010
GOES-13 | GOES-East, 75W | ~12:35pm | 04/14/2010 –
GOES-15 | GOES-West, 135W | ~12:15pm | 12/06/2011 –
Aqua MODIS | LEO | ~1:30pm | 07/04/2002 – 10/31/2012
Hyperion | LEO | ~10:50am | 2011/10/07UTC17:52:08 2012/05/13UTC17:57:19 2012/06/08UTC17:51:58 2012/10/01UTC17:44:40

## Slide 7: Removal of Cloud/Vegetation Pixel

## Slide 8: Long-term stability

## Slide 9: Spatial Uniformity

## Slide 10: Challenges in this Study
Large MODIS viewing angle over the desert
  Viewing angle range for MODIS calibration uncertainty specification
  GOES-East VZN:55.8; GOES-West: 55.6
  MODIS 45 scan angle = about 54 vzn
Difference in GOES and Aqua MODIS observation time
  ~1 – 1.5 hours of observation difference
No collocation with similar viewing/solar viewing geometry
  Challenging for Ray-matching method to transfer the calibration to Aqua/MODIS
Need accurate BRDF model to transfer the GEO trending result to MODIS data
  Assuming that the azimuth angle impact is negligible
MODIS
MODIS
MODIS

## Slide 11: Transferring MODIS to GOES Solar Viewing Geometry
Fixed LEO and GEO viewing geometries
Only the solar viewing geometry affects the BRF
 | GOES-East | GOES-West
Transferring Calibration | 0.58 | 0.72

## Slide 12: TOA reflectance of Sonoran Desert
 | GOES-East | GOES-West
Long-term stability | 34.49(±1.3) | 30.26(±1.11)

## Slide 13: Trending
Wu, 2005, Calcon

## Slide 14: Trending residuals

## Slide 15: Trending Residuals

## Slide 16: Trending Results

## Slide 17: Trending Result
 | Uncertainty
Trending residual | 0.32-0.81

## Slide 18: SBAF Correction with Hyperion Data

## Slide 19: Uncertainty Budget
 | Aqua/MODIS | Target long-term stability* | Transferring calibration* | Trending | Band Correction | Overall Uncertainty
GOES-East | 1.64* | 1.3 | 0.58 | 0.32-0.81 | 0.25-0.34 | 2.21 – 2.34
GOES-West | 1.64* | 1.1 | 0.72 | 0.32-0.81 | 0.25-0.34 | 2.14 - 2.28
*2: uncertainty of target long-term stability can beat down with BRDF correction
*3: Missing correction transfer calibration to account for the solar/viewing azimuth angle difference , especially for GOES-East
*1: MODIS uncertainty specification is applicable to the GOES VZN

## Slide 20: Summary
Use of the sub-area [32.05-32.35N, 114.4-114.7W] as the reference
  avoid of vegeted pixels and artificial feature (high-ways)
Not recommended for near-real-time (operational) calibration
Hyperion data can be used to generate the spectral band adjustment factor
  Relatively spectrally flat over the target and SRF range
Need accurate BRDF model to account for the impact of relative azimuth angle, especially for GOES-East
Absolute uncertainty is about 2-3%

## Slide 21: TOA Reflectance to Dynamic Atm. Components
Chart: Precipitation Water Vapor
Chart: Ozone
| 0.15000000000000019 | 0.15500000000000022 | 0.16 | 0.16500000000000001 | 0.17 | 0.17500000000000004 | 0.18000000000000019 | 0.18500000000000019 | 0.19 | 0.19500000000000001 | 0.2 | 0.20500000000000004 | 0.21000000000000019 | 0.21500000000000019 | 0.22 | 0.22500000000000001 | 0.23 | 0.23500000000000001 | 0.24000000000000019 | 0.24500000000000019 | 0.25 | 0.255 | 0.26 | 0.26500000000000001 | 0.27 | 0.27500000000000002 | 0.28000000000000008 | 0.28500000000000031 | 0.29000000000000031 | 0.29500000000000032 | 0.30000000000000032 | 0.30500000000000038 | 0.31000000000000039 | 0.31500000000000039 | 0.32000000000000045 | 0.32500000000000046 | 0.33000000000000052 | 0.33500000000000052 | 0.34 | 0.34500000000000008 | 0.35000000000000031 | 0.35500000000000032 | 0.36000000000000032 | 0.36500000000000032 | 0.37000000000000038 | 0.37500000000000039 | 0.38000000000000045 | 0.38500000000000045 | 0.39000000000000046 | 0.39500000000000052 | 0.4
G13 | -0.906 | -0.881 | -0.856 | -0.83 | -0.805 | -0.779 | -0.754 | -0.729 | -0.704 | -0.678 | -0.653 | -0.628 | -0.603 | -0.578 | -0.553 | -0.528 | -0.503 | -0.477 | -0.452 | -0.428 | -0.403 | -0.378 | -0.353 | -0.328 | -0.303 | -0.278 | -0.253 | -0.229 | -0.204 | -0.179 | -0.154 | -0.13 | -0.105 | -0.081 | -0.056 | -0.031 | -0.007 | 0.018 | 0.042 | 0.067 | 0.091 | 0.116 | 0.14 | 0.164 | 0.189 | 0.213 | 0.237 | 0.262 | 0.286 | 0.31 | 0.334
G11 | -0.526 | -0.505 | -0.485 | -0.464 | -0.443 | -0.423 | -0.402 | -0.381 | -0.361 | -0.34 | -0.319 | -0.299 | -0.278 | -0.258 | -0.237 | -0.217 | -0.196 | -0.176 | -0.155 | -0.135 | -0.114 | -0.094 | -0.074 | -0.053 | -0.033 | -0.013 | 0.007 | 0.028 | 0.048 | 0.068 | 0.088 | 0.109 | 0.129 | 0.149 | 0.169 | 0.189 | 0.209 | 0.229 | 0.249 | 0.269 | 0.289 | 0.309 | 0.329 | 0.349 | 0.369 | 0.389 | 0.409 | 0.429 | 0.448 | 0.468 | 0.488
Nominal TOA refl: G11=28.7%, G13=29.0%
Chart: AOD
| 0.1 | 0.14000000000000001 | 0.18000000000000013 | 0.22 | 0.26 | 0.30000000000000027 | 0.34 | 0.38000000000000034 | 0.42000000000000026 | 0.46 | 0.5 | 0.54 | 0.58000000000000007 | 0.62000000000000055 | 0.66000000000000081 | 0.70000000000000051 | 0.74000000000000055 | 0.78 | 0.82000000000000051 | 0.86000000000000054 | 0.9 | 0.9400000000000005 | 0.98
G13 | -0.002 | 0.001 | 0.004 | 0.005 | 0.004 | 0.001 | -0.005 | -0.012 | -0.021 | -0.034 | -0.05 | -0.07 | -0.093 | -0.12 | -0.152 | -0.186 | -0.221 | -0.262 | -0.305 | -0.351 | -0.401 | -0.452 | -0.507
G11 | 0.007 | 0.007 | 0.007 | 0.006 | 0.003 | -0.003 | -0.01 | -0.02 | -0.031 | -0.046 | -0.064 | -0.086 | -0.111 | -0.139 | -0.171 | -0.207 | -0.242 | -0.283 | -0.325 | -0.37 | -0.42 | -0.471 | -0.525