## Slide 1: KMA GPRC Report
2014 GSICS Annual Meeting (24-28 Mar 2014) @Darmstadt
Minjin CHOI
Dohyeong KIM, Ho-Seung LEE, Hye-Sook LEE, Tae-Hyeong OH
National Meteorological Satellite Center
Korea Meteorological Administration
- Current Status and Future Plan of KMA -

## Slide 2: CURRENT STATUS OF KMA
Visible channel
  Operate vicarious calibration system combined desert/ocean/cloud targets
  keep studying the moon targets for vicarious calibration
  providing time series of slope and intercept
Infrared channel
  Operate GSICS GEO-LEO Inter-calibration system
  use of two kinds of LEO sensors (MI/COMS & IASI/MetOp-A and MI/COMS & AIRS/Aqua)
  providing time series of mean biases for TBs
Display in website
  Display the method and results of GSICS activities in NMSC homepage
Re-Calibration
  Test an effect of GSICS correction on COMS Level 2 products

## Slide 3: VISIBLE CHANNEL (1/2)
check the change of slope calculated from scatter plot of 4 targets
slope is obtained everyday using data for 30 days ago.
a scatter plot for 30 days data
get a slope/intercept
trend of slope for long period
Ocean
Desert
Water Cloud
DCC
RTM Simulated reflectance
COMS Observed reflectance
2011.04 – 2013.12

## Slide 4: VISIBLE CHANNEL (2/2)
period: 2011.04 – 2013.12
degradation: 5.59% (2.10%/year) from 4 targets on the Earth
3.47% (1.22%/year) from the moon targets
daily slope
from 4 targets
monthly ratio
from moon targets
slope = -0.172
degradation = 5.59%
slope = -0.102
degradation = 3.47%
In spite of relatively short analysis period,
COMS MI visible sensor seems to be maintain its response sensitivity.

## Slide 5: INFRARED CHANNEL (1/3)
check the TB bias for each IR channels using both IASI and AIRS
period: 2011.04 – 2013.12
 |  | BOTH | IASI | AIRS
IR1 | Number | 127948 | 47072 | 80876
 | Bias | 0.2468 | 0.2906 | 0.2212
 | RMSE | 1.5455 | 1.5546 | 1.5402
IR2 | Number | 133908 | 49332 | 84576
 | Bias | 0.2171 | 0.2213 | 0.2146
 | RMSE | 1.4138 | 1.4237 | 1.4079
IR3 (WV) | Number | 66436 | 23775 | 42661
 | Bias | -0.9108 | -0.7471 | -1.002
 | RMSE | 1.2263 | 1.0888 | 1.2965
IR4 (SWIR) | Number | 55038 | 24530 | 30508
 | Bias | 0.1083 | 0.0215 | 0.1780
 | RMSE | 2.8551 | 2.4973 | 3.113
COMS TB (K)
LEO TB (K)
COMS TB (K)
LEO TB (K)
IR1 (10.8μm)
IR2 (12.0μm)
WV (6.75μm)
SWIR (3.75μm)
mean bias 0.2171
RMSE 1.4138
mean bias 0.2468
RMSE 1.5455
mean bias -0.9108
RMSE 1.2263
mean bias 0.1083
RMSE 2.8551
more positive except for IR4
AIRS IASI

## Slide 6: INFRARED CHANNEL (2/3)
COMS TB (K)
TB Bias (K)
COMS TB (K)
TB Bias (K)
trend of TB bias for COMS IR TBs
period: 2011.04 – 2013.12
IR1
IR2
WV
SWIR
large negative bias for warm temperature
different signs
for cold temperature
AIRS IASI BOTH
mean bias 0.2171
RMSE 1.4138
mean bias 0.2468
RMSE 1.5455
mean bias -0.9108
RMSE 1.2263
mean bias 0.1083
RMSE 2.8551

## Slide 7: INFRARED CHANNEL (3/3)
monthly variation of TB bias and RMSE
period: 2011.04 – 2013.12
TB Bias (COMS-LEO)
TB RMSE
IR1
IR2
WV
SWIR
(K)
(K)
AIRS IASI BOTH
mean Bias 0.2171
mean Bias 0.2468
mean Bias -0.9108
mean Bias 0.1083
mean RMSE 1.4138
mean RMSE 1.5455
mean RMSE 1.2263
mean RMSE 2.8551

## Slide 8: WEB DISPLAY
http://gsics.wmo.int/
directly linked!

## Slide 9: WEB DISPLAY
http://nmsc.kma.go.kr/html/homepage/en/gsics/gsicsMain.do
2013, completely construct
web presenting system
(IR and VIS)
Infrared
Visible

## Slide 10: RE-CALIBRATION (1/3) : Concept
To investigate an effect of GSICS calibration coefficient on Level2 products
Test product: Sea Surface Temperature (SST), Solar insolation(INS)
Validation data: buoy around East Asia, pyranometer ground observation of KMA
Period:2012-2013 for SST, 2013 for INS
COMS
Level 1B
TB Reflectance
GSICS coefficients
SST
(IR1, IR2)
INS
(VIS)
raw data
correction
processing & Validation
with buoy
with
pyranometer
WHAT
EFFECT ?

## Slide 11: RE-CALIBRATION (2/3): results of SST
daytime
Before
2012
2013
nighttime
After
Before
After
Comparison of recalibrated SST with observed buoy SST
RMSE 0.86 K
0.81 K
0.67 K
0.81 K
0.81 K
0.94 K
RMSE 0.90 K
0.94 K
Bias 0.50 K
0.08 K
0.01 K
0.01 K
0.01 K
-0.01 K
Bias 0.28 K
-0.02 K

## Slide 12: RE-CALIBRATION (3/3): results of INS
2013
All Sky
After
Before
 | Before |  | After | 
 | RMSE | Bias | RMSE | Bias
All sky | 95.832 | 12.163 | 95.772 | 2.523
cloudy | 91.182 | 21.979 | 89.677 | 2.102
partially | 104.96 | 5.308 | 106.1 | -1.016
COMS INS Algorithm
Comparison of recalibrated INS with observed pyranometer INS
ST : Total solar insolation on ground
SI : Direct solar insolation
SR : Solar insolation by Rayleigh scattering
SA : Solar insolation by Aerosol scattering
Cloud Mask
CF: Cloud Factor
Clear INS
All sky INS
cloudy
clear
effect
improve
in summer season !

## Slide 13: FUTURE PLAN OF KMA
Infrared Channel
  To construct another GSICS operation system (COMS vs. IASI/MetOp-B)
  To further study for large bias of COMS WV channel
  To compare double difference with LEOs (AIRS, IASI)
  To investigate diurnal variation of COMS calibration
Visible Channel
  To improve vicarious calibration method for moon target.
  To refine the operating vicarious system for 4 targets on the Earth
GSICS Webpage
  To keep displaying regular results of GSICS activities like now
RE-Calibration
  To try to test an effect on Clear Sky Radiance

## Slide 14: Thank you Any questions or comments for us?
Notes: Thank you for listening this talk.

## Slide 15: Back-up Slides

## Slide 16: Re-calibration method for SST

## Slide 17: Coefficients (for 2012)
 | a | b
IR1 (10.8㎛) | 1.008 | -2.439
IR2 (12.0㎛) | 1.007 | -1.350
IR3 (6.8㎛) | 1.022 | -5.552
IR4 (3.8㎛) | 1.000 | 0.232
GSICS IR Calibration Coefficient
Used Data
COMS/MI Level 1B data
Period: 2011.04. ~ 2012.03.
Equation: Yλ = aXλ + b
Xλ : COMS TBB before calibration
Yλ : COMS TBB after calibration
a, b : calibration coefficient
λ : wavelength
IR
IR’
SST
SST’
MCSST = C1∙T11㎛ + C2∙(T11㎛-T12㎛) + C3∙(T11㎛-T12㎛) ∙(sec(θ)-1) + C4
 | Day Time |  | Night Time | 
 | C | C’ | C | C’
c1 | 0.985098 | 0.984 | 0.975640 | 0.979
c2 | 2.338343 | 2.547 | 2.496965 | 2.415
c3 | 0.545135 | 0.413 | 0.353631 | 0.484
c4 | -0.321399 | 0.964 | -0.031189 | 2.040
How to recalibrate COMS SST using GSICS IR calibration coefficient

## Slide 18
Channel |  | 2012 before | 2013 before | 2012 after | 2013 after
IR1 channel (10.8 ㎛) | RMSE (K) | 1.3796 | 1.4628 | 1.3079 | 1.4118
 | Bias (K) | 0.3926 | 0.2578 | 0.0913 | -0.0106
 | Slope | 0.9902 | 0.9886 | 0.9981 | 0.9978
 | Intercept | 3.0185 | 3.2856 | 0.0637 | 0.5832
IR2 channel (12.0 ㎛) | RMSE (K) | 1.3185 | 1.3307 | 1.2284 | 1.2892
 | Bias (K) | -0.4485 | 0.1873 | 0.0661 | 0.002
 | Slope | 0.9913 | 0.989 | 0.9983 | 0.9981
 | Intercept | 1.8687 | 3.0963 | 0.5318 | 0.5084
IR1
IR2
Correction effect

## Slide 19: Re-calibration for INS
domain: 30-40°N, 120-130°E, 37 of points
period: 2012.4~2013.3
validation data: AWS INS data in South Korea
GSICS correction coefficient
a : 0.864972 (slope)
b : 0.024749 (intercept)
Solar Zenith Angle
VIS Reflectance
Cloud Factor (0 to 1 range)

## Slide 20: 2013 May DCC web-meeting (R. Bhatt)

## Slide 21: day/night TB bais scatter plot
Using AIRS data
Using IASI data
day (magenta) / night (blue)
period: 2011.04 – 2013.12
IR1 (10.8μm)
IR2 (12.0μm)
WV (6.75μm)
SWIR (3.75μm)
IR1 (10.8μm)
IR2 (12.0μm)
WV (6.75μm)
SWIR (3.75μm)
day: mean bias 0.4042
night: mean bias 0.0911
day: mean bias 0.2488
night: mean bias 0.1876
day: mean bias -0.7731
night: mean bias -1.1688
night: mean bias 0.178
day: mean bias 0.3909
night: mean bias 0.2052
day: mean bias 0.2555
night: mean bias 0.1886
day: mean bias -0.7322
night: mean bias -0.7615
night: mean bias 0.0215

## Slide 22: the number of collocation pixels
period: 2011.04 – 2013.12
sum of IASI and AIRS collocation pixels
monthly values
IR1 IR2 IR3 IR4