## Slide 1: NOAA/NESDIS/STAR GOES/GOES-R Integrated Cal/Val System (ICVS) Long-Term Monitoring (LTM) Development
Fangfang Yu, Fuzhong Weng, Xiangqian Wu, and Ninghai Sun
Satellite Meteorology and Climatology Division
Center for Satellite Applications and Research
National Environmental Satellite, Data and Information Service

## Slide 2: NOAA/NESDIS/STAR ICVS LTM
S-NPP: ATMS/CrIS/VIIRS/OMPS
NOAA/Metop: AVHRR/HIRS/AMSU/MHS
DMSP …
GOES : Imager/Sounder
Spacecraft, space weather instruments
GOES-R: ABI (after launch) … under development
LEO (NPP/JPSS, NOAA/Metop …)
GEO(GOES/GOES-R)

## Slide 3: GOES/GOES-R ICVS Objectives
Provide instrument scientists with calibration healthy status
Provide users with information of radiance quality for product generation.
  Users’ Request in the 2013 Users Workshop: Satellite operational anomaly alerting system for the GOES-R series
Facilitate communication of instrument anomalies amongst instrument engineers, calibration scientists and user community.

## Slide 4: Web-based System…
Focus on long-term monitoring of real-time on-orbit instrument performance and sensor data quality
Spacecraft performance monitoring
Space weather instrument performance monitoring
Weather instrument Monitoring
(evolved from GOES IPM system)
Telemetry
Calibration Coefficients
Detector Noise
Radiance Quality (GSICS)
Under development
Upgrading

## Slide 5: GOES Imager/Sounder ICVS Monitoring
Imager and Sounder ICVS monitoring currently available at: http://www.star.nesdis.noaa.gov/icvs/

## Slide 6
SPLK
TELE
CAL
Noise
Plotting and Report
Extract individual Sounder parameter and calculate the statistics
SPLK
BB
Data Subset
Calibration related dataset
Imager monitoring
Sounder monitoring
BB
Extract the individual Imager parameter and calculate the statistics
TELE
Noise
Plotting and Report
CAL
Near Real-time L1B data
GOES ICVS Daily Data Processing Flowchart

## Slide 7: Current GOES ICVS LTM Parameters
Satellite | Instrument | Calibration Related Telemetry Parameters | Calibration Target Signals | IR Calibration Coefficients | Detector Noise
GOES-13 | Imager | 14 | 4x5channels (2-8 det/chan) | 2x4 channels (2 det/chan) | 2x5chan (2-8 det/chan)
 | Sounder | 16 | 2x19channels (4 det/chan) | 2x19 channels (4 det/chan) | 2x19 channels (4 det/chan)
GOES-15 | Imager | 14 | 4x5 channels (2-8 det/chan) | 2x4 channels (2 det/chan) | 2x5 channels (2-8 detectors)
 | Sounder | 16 | 2x19 channels (4 det/chan) | 2x18 channels (4 det/chan) | 2x19 channels (4 det/chan)

## Slide 8
Multiple panels for a variety of temporal scales
Diurnal variation for different PRT detectors
  Inhomogeneous heating of BB
  midnight heating and the latency
Seasonal Variation
  strongest after-midnight heating around the summer solstice
Long-term trending
  more apparent at certain time

## Slide 9
Nominal calibration coefficients
diurnal variation
  Erroneous Midnight calibration coefficients
Seasonal variation
  erroneous midnight calibration coefficients increase as approaching summer
  jump with patch temperature switch
Long-term trending
  - Increasing absolute coefficients values as sensor degrades

## Slide 10
MBCC slopes cause jumps here
Clamp pos
Due to contaminated splk

## Slide 11
Noise jumped suddenly in early July 2012
Data outage from 09/12/2012-10/18/2012 , outgassing for Sounder
Data outage 05/22/2013- 06/06/2013
Data outage 05/22/2013- 06/06/2013

## Slide 12: Examples of using GOES ICVS (GOES IPM) to investigate the root causes of calibration anomalies

## Slide 13: GOES-15 Imager Calibration Anomaly in March 2012
First revealed by ECMWF monitoring of Tb
Immediately identified with GOES-IPM monitoring from the change of calibration coefficients. Results of GOES-IPM system helped instrument engineers focus on investigating the final root cause in the ground system – change of SPS version
Evaluated with the GSICS GEO-LEO and GEO-GEO bias monitoring
Wu et al. 2012
Notes: http://cimss.ssec.wisc.edu/goes/blog/archives/9995

## Slide 14: GOES-15 SNDR Ch 1-5 Anomaly in EMC GDAS Operational Monitoring …
http://www.emc.ncep.noaa.gov/gmb/gdas/radiance/esafford/wopr/index.html

## Slide 15
GOES-15 SNDR Ch 1-5 Anomaly in EMC GDAS Operational Monitoring …

## Slide 16: Conclusions and Path Forward
A very powerful tool
  Instrument performance health monitoring
  Diagnose the anomalies
  Address the users questions or concerns on the calibration changes
  Evaluate/verify the instrument performance
    post-launch test period
    recovery period right after instrument gets re-activated.
Continue to evolve the GOES ICVS for GOES-R ABI monitoring purpose
  Implement and enrich the data quality monitoring
    Inter-calibration/inter-comparison
  Implement and improve the calibration alert notification function
  Provide flexible customized datasets for analysis