## Slide 1: Calibrating raw data of UVN instrumentation – Lessons learned for a UVN GSICS sub-group on instrument characterisation
Rűdiger Lang, Rose Munro, Antoine Lacan, Christian Retscher, Gabriele Poli, Michael Grzegorski , Andriy Holdak, Rasmus Lindstrot, Ed Trollope

## Slide 2
Outline
On-ground and In-flight UVN characterisation
Some examples from the GOME-2 on-ground and in-flight characterisation efforts (and the lessons learned)
Generic recommendation for UVN characterisation from GOME-2 lessons learned: why can’t we go down to 1% or below in the lab?
Summary from the 1st EO workshop on UVN characterisation (ESRIN, June 2013)
Summary and potential work-outline for a GSICS UVN sub-group on instrument characterisation

## Slide 3: Goniometry: Angular dependence for Solar Measurements (AIRR) Comparison between current key-data and high res. elevation angle grid
Key-data – angle resolution
In-flight derived values with increased angular resolution
Band 3 solar signals at 420 nm
raw
calibrated
key-data
Solar Elevation angle
Solar Elevation angle
GOME-2/Metop-A
GOME-2 FM3 Metop-A

## Slide 4: Goniometry: Angular dependence for Solar Measurements (AIRR) Comparison between current key-data and high res. elevation angle grid
FM3 on-ground
key-data
In-flight data
The Angular dependence of irradiance on the diffuser in elevation and solar azimuth I0(f,e) is
difficult to measure on-ground (long-measurement period / in vacuum)
but one can derive it from in-flight data
1 year of in-flight data needed
Of relevance also for other current and future missions!
GOME-2 FM3 Metop-A

## Slide 5
SAA: 0.5 resolution (317 to 333)
Elevation: 0.1 resolution (-1.5 to 1.5)
Similar to L1-processor angle fine-grid as defined in initialisation file!
Goniometry: Improved AIRR measurements for FM-2 TNO FM2 delta calibration campaign (Kenter et al.)
FM2 Metop B
Higher angular resolution
FM3 Metop A
Original angular resolution
GOME-2 FM3 Metop-A
GOME-2 FM2 Metop-B

## Slide 6: Goniometry: Angular dependence for Solar Measurements (AIRR) Comparison between current key-data and in-orbit derived data Angular domain – azimuth angle at 311 nm
In-flight data
FM3 on-ground key-data
FM3 in-flight data
The Angular dependence of irradiance on the diffuser in elevation and solar azimuth I0(f,e) is
difficult to measure on-ground (long-measurement period / in vacuum)...
...and can lead to artefacts in the long-term time-series of reflectances!
GOME-2 FM3 Metop-A

## Slide 7: Goniometry: Angular dependence for Solar Measurements (AIRR) Comparison between current key-data and in-orbit derived data Spectral domain – Channel 1
The Angular dependence of irradiance on the diffuser in elevation and solar azimuth I0(f,e) is
difficult to measure on-ground (long-measurement period / in vacuum)...
...and can lead to artefacts in the spectral domain of the solar mean reference measurements!
GOME-2 FM3 Metop-A
FM3 on-ground key-data
FM3 in-flight data

## Slide 8
Findings during GOME202-2 Delta Calibration Radiometric Calibration – On-ground Irradiance response
Irradiance Calibration
Large deviations
+/- 4-5% !!!
Radiometric Calibration Investigations
Same principle setup but measured at different times (i.e. With new realignments)

## Slide 9: Findings in-orbit after GOME202-2 Delta Calibration Channel 3: Key-data artifacts
“Zeta” in channel 3
(pol. Sensitivity for 45 degrees polarised light)
FFT filter removes small scale structure
GOME-2 FM2 Metop-B

## Slide 10: Findings in-orbit after GOME202-2 Delta Calibration Channel 3: Key-data artifacts – light sources used during calibration
Origin of small scale structure:
Xe calibration lamp
Residuals are the result of an instability of the lamp during measurement period
(otherwise the spectral structures are divided out)
GOME-2 FM2 Metop-B
Courtesy G. Otter, TNO/TPD

## Slide 11: Findings in-orbit after GOME202-2 Delta Calibration Channel 3: Key-data artifacts – viewing angle direction
Angular dependence of “Zeta” in channel 3 (Chi for Zeta)
(pol. Sensitivity for 45 degrees polarised light)
GOME-2 FM2 Metop-B
Key-data has been corrected in recent updates for FM2 and FM3
(May/June 2013)!

## Slide 12
Residuals are small but spectrally persistent and varying in viewing angle
GOME-2 FM2 Metop-B
Key-data has been corrected in recent updates for FM2 and FM3
(May/June 2013)!
Findings in-orbit after GOME202-2 Delta Calibration Channel 3: Key-data artifacts – impact on radiances

## Slide 13: Findings in-orbit after GOME202-2 Delta Calibration Channel 3: Key-data artifacts – impact on retrievals
Initial retrieval differences observed (preliminary)
NO2 retrieval:
425 to 497 nm
Courtesy:
I. deSmet, BIRA
Old key-data
New key-data
TNO2
RMS
TNO2
RMS
GOME-2 FM2 Metop-B

## Slide 14
Metop-A/B GOME-2 Radiometric accuracy and calibration
Earthshine measurement co-location FM2 to FM3 / Co-location criteria
Co-location criteria:
Area overlap > 80%
Geo. AMF diff < 2%
Relative O2 A-Band residual (ROR) smaller than empirical threshold
GOME-2 FM3 Metop-A
GOME-2 FM2 Metop-B
Metop-A (M02): FM3
Metop-B (M01): FM2
Time difference:
49 min (half an orbit)

## Slide 15
26th December 2012
10th January 2013
Metop-A/B GOME-2 Radiometric accuracy and calibration
Earthshine co-located Metop-A/B measurements – co-locations
GOME-2 FM3 Metop-A
GOME-2 FM2 Metop-B

## Slide 16
Metop-B / FM2 Radiometric accuracy and calibration
Earthshine measurement co-location to Metop-A / FM3 / Radiances
Radiance residuals are dominated by FM3 degradation signature
GOME-2 FM3 Metop-A
GOME-2 FM2 Metop-B

## Slide 17
Example for an average over multiple FM2/FM3 residuals in reflectivity.
Prelim. Approach:
A linear background is subtracted per channel to account for the difference in reflectivity-degradation and for broad-band differences in the different observed atmospheric paths per instrument.
Looking for the small scale structures per channel
GOME-2 FM3 Metop-A
GOME-2 FM2 Metop-B
Metop-A/B GOME-2 Radiometric accuracy and calibration
Earthshine co-located Metop-A/B residuals
Background predominantly from Metop-A degradation

## Slide 18
Residual in reflectivity
+ Lin. background subtracted per channel
+ 20 pix. moving average smoothing
+ Relative Oxygen-A band Residual (ROR) selection criterium
Metop-A/B GOME-2 Radiometric accuracy and calibration
Earthshine co-located Metop-A/B reflectances and residuals
GOME-2 FM3 Metop-A
GOME-2 FM2 Metop-B

## Slide 19
GOME-2 FM3 Metop-A
GOME-2 FM2 Metop-B
FM2-East/FM3-West
FM2-West/FM3-East
Nadir
Separation between East/West/Nadir at +- 20 degree viewing-angle
Metop-A/B GOME-2 Radiometric accuracy and calibration
Earthshine co-located Metop-A/B residuals / average background subtracted

## Slide 20
GOME-2 FM3 Metop-A
GOME-2 FM2 Metop-B
Metop-A/B GOME-2 Radiometric accuracy and calibration
Earthshine co-located Metop-A/B residuals / Background-subtracted
plus degradation residual from Metop-A
Background subtracted co-located Metop-A/B residual
Background subtracted temporal degradation residual Metop-A

## Slide 21
GOME-2 FM3 Metop-A
GOME-2 FM2 Metop-B
Background subtracted co-located Metop-A/B residual
without the temporal degradation contribution of Metop-A
Metop-A/B GOME-2 Radiometric accuracy and calibration
Earthshine co-located Metop-A/B residuals / Background-subtracted
plus degradation residual from Metop-A subtracted

## Slide 22
Recommendations for On-Ground Characterization
Time Time Time
Considerations for On-Ground Characterisation Campaign (I)
Time … Time … Time
  Characterisation campaigns should be long enough to allow for measurements to be repeated for consistency checking
  Assess impact of lamp position & alignment errors
  Allow sufficient time for stabilisation

## Slide 23
Considerations for On-Ground Characterisation Campaign (I)
Environment
  Essential that all characterisation measurements are carried out in thermal vacuum and that the thermal environment including gradients is representative of the in-orbit situation
  Scan-angle dependencies should also be characterized in vacuum
Recommendations for On-Ground Characterization
Environment

## Slide 24
Considerations for Measurements
Procedures I
  Alignment procedures should be documented and reproducible with photo/video documentation
  Ensure reproducibility of distance measurements
  Close attention should be paid to frames of reference, coordinate systems & angles (e.g. GOME-2 flip of elevation angles diagnosed in orbit)
  All sources should be well commissioned prior to the start of measurement
  Radiometric calibration must be connected to standards e.g. NIST
  All measurements and procedures must be traceable & under configuration control including software versions and documentation of which precise measurement is used in the generation of key data
  Data processing should be automated as far as possible
Recommendations for On-Ground Characterization
Procedures I

## Slide 25
Considerations for Measurements
Procedures II
  Check temperature sensitivity in case thermal stability is not as expected
  Ensure sufficient angular discretisation for characterisation of diffuser BSDF
  For slit function characterisation need requirements from the data analysis activity (signal:noise, spectral coverage and sampling, source commissioning requirements etc) before planning the measurements
  Ensure that all required supplementary calibration measurements (e.g. dark signal etc) are taken close to the time of each measurement (GOME-2 monitoring block)
  Ensure sufficient sampling points for straylight characterisation
  Ensure that the slit is overfilled
Recommendations for On-Ground Characterization
Procedures II

## Slide 26
Considerations for In-Orbit Calibration & Performance Verification
  Consider taking necessary in-orbit calibration measurements adjacent to Sun measurements if possible
  Take dark measurements under the same conditions as measurements
  Much longer measurement time will be needed for e.g. WLS over diffuser as opposed to WLS direct (if used for monitoring diffuser) so it is necessary to ensure this is compatible with lamp lifetime and recommended use
  Take many monitoring and calibration measurements early in instrument life
  Be wary of coatings etc that will require time to stabilise and outgas
  Be aware of any temperature dependencies of output or aging issues in on-baord targets (e.g. LED/WLS).
  …
Recommendations for In-Orbit Calibration
Summary

## Slide 27
  Instrument should typically be stored in a container over pressured using nitrogen
  Regular reactivation is required
  Extreme attention to cleanliness required
  Assess what monitoring measurements can be made during regular re-activation in ambient.
  …
Recommendations for In-Orbit Calibration
Summary cnt.

## Slide 28
From level-2 perspective (Diego Loyola, DLR):
Improvements in fit qualities for direct fitting approaches (as with respect to standard DOAS) with however higher sensitivity to level-1 calibration issues call for inter-calibrated UVN level-1 products
SCIAMACHY/GOME-1 experience (Thijs Krijger, SRON)
Better documentation and availability of all on-ground measurements (also those not used in the “official” data-packs, for recovery or development of in-orbit error or degradation mitigating methods (“dirty mirror” - model).
On-board and on-ground reference sources (N. Fox, NPL)
Demands rigorous instrument model during calibration with all potential error sources known and documented. He suggests to use tuneable lasers for Stray-light.
He claims it is possible to design an on-board reference source to accuracies significantly below 1%, which is very insensitive to the degradation of the reference detector and which in turn would be linked to a different source which does the spectral changes. This approach would be very insensitive to the degradation of the reference detector (Fox et al, Phil. Trans. R. Soc. A 2011 369, 4028-4063).
Notes from the 1st EO lessons learned meeting ESRIN
June 2013 – 1/3

## Slide 29
From MERIS/OLCI diffuser and stray-light (Ludovic Bourg, ACRI-ST, Juergen Fischer FU Berlin):
MERIS sees hardly any diffuser degradation (below 4% full lifetime at 400nm).
However much more detailed stray-light model was needed. Final stray-light level correction term is also proportional to the observed brightness with a quadratic factor in pixel space.
The initially provided on-ground (very complex) ray-tracing model correction levels for the instrument stray-light was not sufficient, although it already accounts for a large proportion of the stray-light contribution.
POLDER-3/Parasol experience (Patrice Henry)
A posteriori developed degradation model available.
After correction for degradation they then apply pre-dominantly vicarious calibration methods (“Muscle and Muse” analysis suite including database containing reference data per target).
Achieved accuracies for vicarious calibration are not better than 1%. The latter can only be achieved by combining different methods and sources.
Notes from the 1st EO lessons learned meeting ESRIN
June 2013 – 2/3

## Slide 30
From MERIS/OLCI diffuser and stray-light (L. Bourg, ACRI-ST, J. Fischer FU Berlin):
MERIS sees hardly any diffuser degradation (below 4% full lifetime at 400nm).
However much more detailed stray-light model was needed. Final stray-light level correction term is also proportional to the observed brightness with a quadratic factor in pixel space.
The initially provided on-ground (very complex) ray-tracing model correction levels for the instrument stray-light was not sufficient, although it already accounts for a large proportion of the stray-light contribution.
OMI/Sentinel5p (Quintus Kleipool/Marcel Dobber, KNMI)
Not having a scan-mirror is an advantage and the crucial (OMI) mirror is very well hidden inside the instrument.
On-ground calibration has missing angular dependence measurements in vacuum and not enough for operational range of conditions.
Radiation damage has triggered a large increase of dead pixel, increase of dark current, and RTS changes for OMI even though a lot of additional shielding has been added.
WLS on OMI is difficult to use for non-linearity because it is not stable enough
Notes from the 1st EO lessons learned meeting ESRIN
June 2013 - 3/3

## Slide 31
On-ground calibration efforts need to be improved at all levels: time, documentation, availability, and accuracy of data and the related error budgets.
Vicarious calibration can complement these efforts but so far accuracy levels below 1% can not be achieved.
Below 1% radiometric accuracy levels are difficult to achieve if achievable at all through on-ground calibration alone. On-board calibration and vicarious calibration methods need to be added.
Huge amount of effort has been invested to compensate for avoidable deficiencies in on-ground characterisation for past or currently flying UVN instrumentation.
Some proposals are on the table to significantly (fundamentally) improve our capabilities for accurate on-ground characterisation (e.g. Fox et al.)
1st EO lessons learned meeting in ESRIN, June 2013 called for detail recommendation on “best practices” for on-ground calibration in-line and supported by the CEOS activities of best-practices for calibration.
Summary
Best practices for UVN on-ground and in-flight calibration and characterization
Shall and can we support these activities with UVN GSICS activities?

## Slide 32
Any suggestions/ideas/comments?

## Slide 33
Left up: Residual Metop-B with V-LIDORT
Residual Metop-A with V-LIDORT
Bottom: Residuals between M-A and M-B
Broad-band structure in M-A channel 3
Small-scale structures in M-B channel 3
GOME-2 FM3 Metop-A
GOME-2 FM2 Metop-B
Metop-A/B GOME-2 Radiometric accuracy and calibration
Earthshine co-located Metop-A/B residuals /
Interpretation with forward model – V-LIDORT