## Slide 1: KMA Agency Report
Hyesook Lee
Dohyeong Kim, Tae-Hyeong Oh, Minju Gu
National Meteorological Satellite Center
Korea Meteorological Administration

## Slide 2: CURRENT STATUS OF KMA
Visible channel
  Operate vicarious calibration system using targets on the Earth
  keep studying the moon targets for vicarious calibration
Infrared channel
  Operate GSICS GEO-LEO Inter-calibration system
   MI/COMS & IASI-A/-B and MI/COMS & AIRS/Aqua
  Keep studying the cold bias correction in SW with SRF correction
Preparation for the Demonstration Phase
  Preparing COMS IR to enter the Demonstration Phase of the GSICS Procedure for Product Acceptance

## Slide 3: VISIBLE CHANNEL(1/2)
Time series of slopes and intercepts
The degradation is about 5.67%(1.67%/year) from Apr.2011 to Sep. 2014.
a scatter plot for 30 days data
get a slope/intercept
trend of slope for long period
3

## Slide 4: VISIBLE CHANNEL(2/2)
Visible Calibration by Moon target
  In 79 cases of COMS lunar observation
  Compare the result of IMPS and GIRO_fourth realese
    Degradation(IMPS) is about -5.97%(-1.41%/year)
    Degradation(GIRO) is about -4.26%(-0.98%/year)
Issues
  Vary the ratio of observed irradiance and ROLO irradiance through the threshold
  And also change the ratio depend on the Moon area size
  Need more test to define the Moon pixel and Moon area size

## Slide 5: INFRARED CHANNEL(1/3)
5
check the TB bias for each IR channels using IASI_A, IASI_B and AIRS
period: 2011.04 – 2014.12(IASI_A, AIRS), 2013.08 – 2014.12(IASI_B)
AIRS IASI_A IASI_B
 |  | TOTAL | IASI_A | IASI_B | AIRS
IR1 | Number | 207072 | 71524 | 47349 | 88199
 | Bias | 0.2694 | 0.2859 | 0.3296 | 0.2237
 | RMSE | 1.5041 | 1.5024 | 1.4581 | 1.5295
IR2 | Number | 217960 | 75428 | 50108 | 92424
 | Bias | 0.2312 | 0.2186 | 0.2753 | 0.2175
 | RMSE | 1.3806 | 1.3781 | 1.3519 | 1.3979
IR3 (WV) | Number | 108183 | 36696 | 24648 | 46839
 | Bias | -0.8614 | -0.7485 | -0.7652 | -1.0005
 | RMSE | 1.1780 | 1.0805 | 1.0862 | 1.2920
IR4 (SWIR) | Number | 94528 | 37490 | 24304 | 32734
 | Bias | 0.0905 | 0.0374 | 0.0506 | 0.1811
 | RMSE | 2.6635 | 2.4108 | 2.4072 | 3.0869
RMSE 1.5041
RMSE 1.3806
RMSE 1.1780
RMSE 2.6635
IR1 (10.8μm)
IR2 (12.0μm)
WV (6.75μm)
SWIR (3.75μm)
LEO TB (K)
COMS TB (K)
COMS TB (K)
LEO TB (K)

## Slide 6: INFRARED CHANNEL(2/3)
IR1
IR2
WV
SWIR
trend of TB bias for COMS IR TBs
period: 2011.04 – 2014.12(IASI_A, AIRS), 2013.08 – 2014.12(IASI_B)
AIRS IASI_A IASI_B TOTAL
TB Bias (K)
TB Bias (K)
COMS TB (K)
COMS TB (K)
Bias 0.2237 0.2859 0.3296 0.2694
Bias 0.2175 0.2186 0.2753 0.2312
Bias -1.0005 -0.7485 -0.7652 -0.8614
Bias 0.1811 0.0374 0.0506 0.0905
IR2
RMSE 1.5041
RMSE 1.3806
RMSE 1.1780
RMSE 2.6635

## Slide 7: INFRARED CHANNEL(3/3)
7
TB Bias (COMS-LEO)
TB RMSE
IR1
IR2
monthly variation of TB bias and RMSE
period: 2011.04 – 2014.12(IASI_A, AIRS), 2013.08 – 2014.12(IASI_B)
AIRS IASI_A IASI_B TOTAL
(K)
(K)
RMSE 1.5295 1.5024 1.4581 1.5041
SWIR
Bias 0.2237 0.2859 0.3296 0.2694
Bias 0.2175 0.2186 0.2753 0.2312
Bias -1.0005 -0.7485 -0.7652 -0.8614
Bias 0.1811 0.0374 0.0506 0.0905
RMSE 3.0869 2.4108 2.4072 2.6635
RMSE 1.3979 1.3781 1.3519 1.3806
RMSE 1.2920 1.0805 1.0862 1.1780
WV

## Slide 8: Cold Bias Correction in SW
TB bias btw. MI and IASI reduced by ~ 0.4K(-0.71K  -0.32K)

## Slide 9: COMS IR for Demo Phase
Preparation for COMS IR to enter the Demonstration Phase of the GSICS Procedure for Product Acceptance(GPPA)
  Preliminary version of ATBD and sample data (Nov. 2014)
  Checking sample data format

## Slide 10
Thank you