## Slide 1: Extending DCC to other bands and DCC ray-matching
Doelling and input from many others, March 18, 2015

## Slide 2: Overview
DCC application to other visible bands
Applicability of DCC to derive scan angle dependency
DCC ray-matching to validate the DCC mode reflectance or absolute calibration

## Slide 3: Extending DCC calibration to other bands
Apply the MODIS 0.65µm DCC calibration to other bands
  0.47µm, 0.55µm, 1.24µm, 1.37µm, 2.12µm bands
  Terra and Aqua MODIS Collection 6
  Plots in Terra to Aqua scaling paper (Doelling et al. 2015)
Apply the MODIS 0.65µm DCC calibration to VIIRS
  Mbands, 0.48µm, 0.55µm, 0.65µm, 0.86µm, 1.6µm, 2.25µm
  I bands, 0.65µm, 1.6µm
  VIIRS NASA LandPeate calibration applied
  Extend timeline to end of 2014 (Bhatt et al. 2014)
Apply to Met-9 0.86µm and 1.6µm bands

## Slide 4: Terra and Aqua MODIS C6 DCC

## Slide 5: Terra and Aqua MODIS C6 DCC

## Slide 6: NPP-VIIRS DCC and Libya-4

## Slide 7: NPP-VIIRS DCC and Libya-4

## Slide 8: DCC calibration standard errors
Band | Terra | Aqua | NPP-M | NPP-I
0.47 µm | 0.6 | 0.6 | 0.7 | 
0.55 µm | 0.7 | 0.6 | 0.7 | 
0.65 µm | 0.6 | 0.6 | 0.6 | 0.6
0.86µm |  |  | 0.45 | 
1.24 µm | 0.6 | 0.9 |  | 
1.37 µm | 1.0 | 1.3 |  | 
1.6 µm |  |  | 2.4 | 2.3
2.2 µm | 1.8 | 3.4 | 1.9 | 
• DCC has a standard error about the fit of < 0.7%, this includes instrument noise for wavelengths < 1.2 µm
• For wavelengths > 1.1 it is band specific
• Number of years needed to detect a trend the magnitude of 2s at 95% confidence at 50% probability is ~10 years
• For DCC the number year to detect a trend > 2s is ~5 years

## Slide 9: GEO examples
0.65µm
s=0.9%
0.86µm
s=0.6%
1.6µm
s=4.7%
It seems that DCC works best for the 0.86µm band, verified by both Met-9 and VIIRS
Based on conversation with Sebastien, you can identify in the 0.65µm and 1.6µm the calibration discontinuity in late 2011.
s~2% 2007-2011

## Slide 10: MODIS RVS calibration using DCC
MODIS uses a mirror to scan the earth.
The mirror will have differences in reflectivity that change over time
The response versus scan angle (RVS) was characterized at prior to launch and changes over time on orbit
Currently the MODIS team uses deserts trends referenced at launch over time to detect trends in various RVS positions
Can DCC be used to track RVS dependencies?
Every band will have a different RVS correction

## Slide 11: MODIS scan mirror schematic
The Solar Diffuser (SD) and Lunar scan angles can be perfectly calibrated, however, other scan angles may not be.
Apply DCC algorithm in segments of 100 frame numbers
Currently Aqua-MODIS Band 1 (0.65µm) shows a 1.53% degradation

## Slide 12: Aqua-MODIS C6 0.65µm DCC calibration by frame number
RVS trends may not be linear
FRAMES 1-150
With lunar
FRAMES 1-150
With SD

## Slide 13
F# | 75 | 200 | 300 | 400 | 500 | 600 | 700 | 800 | 900 | 1000 | 1100 | 1200
trend | -1.9 | -2.2 | -2.1 | -2.1 | -1.6 | -1.2 | -1.4 | -1.4 | -0.9 | -0.7 | -0.8 | -0.9
s | 1.0 | 1.0 | 1.0 | 1.1 | 0.9 | 0.9 | 0.8 | 0.8 | 0.8 | 0.7 | 0.7 | 0.7
Aqua-MODIS 0.65µm frame number DCC trends
• The SD frame numbers have the lowest trends
• Working with MODIS group to remove the scan angle dependencies
• DCC have the potential to derive scan mirror dependencies

## Slide 14: DCC ray-matching
Verify the Aqua-MODIS BRDF corrected DCC multi-year mode radiance over the GEO domain
  Assume the same DCC are captured by both Aqua-MODIS during 1:30 PM overpass time and GEO, although not exactly coincident and by the same viewing conditions
  Rely on DCC BRDF model to remove sampling dependencies
DCC ray-matching vs gridded ray-matching
  DCC has lowest SBAF uncertainty, less stringent angle, temperature matching, but are DCC available over the matching domain
  Gridded has large SBAF uncertainty, stringent angle matching, and use all scenes
  Both should provide the consistent calibration
This calibration can then be used to validate the DCC mode reflectance

## Slide 15
Jan 1, 2011, MTSAT2
Grid VS
Grid IR
DCC az<15
DCC az<30

## Slide 16: MTSAT-2/Aqua-MODIS raymatching using all 0.5° lat/lon regions
SBAF based on radiance applied

## Slide 17: DCC Ray-matching
1° lat/lon grid, MTSAT-2, July 20, 2011 2:32 GMT
92-km
302-km

## Slide 18: MTSAT-2 10-km DCC ray-matching
Monthly regression 10-km DCC ray matching
Daily gains 10-km DCC ray matching
Adding DCC BRDF to remove angular matching differences increased the noise

## Slide 19: Gridded and DCC ray-matching
Investigating very carefully, SBAF, angular difference impact, sensor linearity, and Met-7 sub-satellite point. MODIS scan angle dependency unresolved

## Slide 20: Gridded and DCC ray-matching
GEO s(%) | Gridded | 10-km | 30-km
GOES-13 | 0.88 | 0.44 | 0.42
GOES-15 | 0.86 | 0.42 | 0.35
MET-9 (0.75 lunar) | 0.83 | 0.41 | 0.43
MET-7 | 0.85 | 0.55 | 0.60
MTSAT-2 | 1.12 | 0.45 | 0.39
GEO bias(%) | Gridded | 10-km | 30-km
GOES-13 | 0.19 | 0.11 | 
GOES-15 | 0.27 | 0.10 | 
MET-9 | -0.16 | 0.09 | 
MET-7 | -0.31 | 0.00 | 
MTSAT-2 | -1.07 | 0.07 | 

## Slide 21: Future DCC ray-matching and DCC absolute calibration
Determine issue with gridded and DCC ray-matching
Compare the DCC ray-matched calibration with the DCC mode reflectance