## Slide 1: Combining Vicarious Calibrations
F. Yu and X. Wu
NOAA GPRC

## Slide 2: Outlines
Research of combined method at NOAA
Individual Vicarious Calibrations (GOES-12)
  Desert
  Ray-matching
  DCC
  Lunar
  Combined Vicarious Calibration
Sensitivity analysis
  Desert + ray-matching + DCC + lunar
  DCC + lunar
Absolute calibration
Conclusions

## Slide 3: background
Combined Vicarious Calibration

## Slide 4: 2010 Toulouse
Reviewed and compared several vicarious calibrations.
Combined all results by normalizing to a common Day 1.
Considered seasonal variations (even for DCC & star)
  Shape, phase, amplitude
  May penalize the good data

## Slide 5: 2011 Daejeon
Applied the recursive filter to GOES-12 star data
  Normalized to a common intensity;
  Applied the filter to daily data;
  Computed the monthly value (comparable to DCC).

## Slide 6: 2014 Darmstadt
Removed target seasonal variation if can be characterized (e.g. desert).
Used monthly statistics for DCC, ray-matching, star, and desert data.
Normalized the monthly data to the estimated Day1 obs.
Applied the recursive filter.
Rename it as integrated method.

## Slide 7: GOES-12 AS EXAMPLE
Combined Vicarious Calibration

## Slide 8: GOES Imager Visible Calibration Methods
Reference targets:
  Sonoran desert
  Deep Convective Cloud (DCC)
  Ray-matching (e.g. over cloud)
  Moon
  Stars
Absolute calibration accuracy was achieved by calibrating the GOES Imager visible data traceable to Aqua MODIS Band 1 C6 standard
Absolute calibration
Absolute/relative calibration, depending on SBAF availability
Relative calibration

## Slide 9: Sonoran Desert
Daily obs.
Remove the cloud contaminate daily obs. , then averaged the clear-sky data at monthly interval
Remove the seasonal variations caused by BRDF and atm. variations
Yu, F., Wu, X., Grotenhuis, M. & Qian, H. (2014). Inter-calibration of GOES Imager visible channels over the Sonoran desert, J. Geo. Res.-Atmosphere, 119, 8639-8659, doi:10.1002/2013JD020702.

## Slide 10: GOES/MODIS Ray-Matching
Ratio of ray-matched GOES and Aqua/MODIS reflectances
  < 10 degree from GEO nadir
  < 1% difference in viewing angles
  < 10 minutes in viewing times
  < 1km in central pixels
Select uniform cloudy scenes
  MODIS reflectance > 50%
  GOES reflectance > 25%
  CoV < 3% within the 3x3 array
  Monthly sample size > 5000 to ensure robust statistics.
  Analyze the time-series of monthly mean reflectance ratios.

## Slide 11: DCC Reflectance
Followed the GSICS DCC ATBD
Monthly sample size > 2000 to ensure robust statistics
Used the median value of monthly DCC reflectance. (Mode and mean were considered.)

## Slide 12: Lunar Irradiance Ratio
Time-series of GOES/ROLO lunar irradiance ratio.
Almost symmetric distribution of lunar phase angles during GOES-12 mission life
More observations since Nov 2005 when lunar obs. were scheduled.

## Slide 13: COMBINED VICARIOUS CALIBRATION
Combined Vicarious Calibration

## Slide 14: Theory
Why does target measurements vary in time?
  Degradation/Anomaly of the instrument
  Natural variability of the target
  Inherent uncertainty of vicarious calibration
Combining multiple vicarious calibrations
  Active: Experts assigns weight (can be binary) to each method / data based on a priori knowledge, e.g., B. Fougnie, GSICS web meeting, Oct. 22, 2014.
  Passive: “1” is common to all methods. “2” & “3” vary among methods but tend to be mutually independent (this leads to iid).
    Weighted by STDV.
    Recursive filter
Use any / all where appropriate.

## Slide 15: NOAA Implementation
Normalized to a common Day1 reflectance
Trending fitting
Monthly Statistics
Combining the Normalized Data
Removal of predictable oscillations

## Slide 16: Multiple Vicarious Calibrations
Trending of the Normalized Results for Each Method
Collection of Normalized Observations
Comparable relative calibration accuracy.
Similar degradation patterns over different reference targets indicate that the impact of spectral response function shift, if any, is very small and negligible on the reflectance.

## Slide 17: Recursive Filter Method
Adapted (with minor mod) from Wu et al. 2010.
Assumption: all and only outliers are not normally distributed.
Iteratively identify and remove outliers until the remaining data are normally distributed.
  for iteration=1, max_iteration do begin ; max_iteration = 4
    result = regress (x, y, yfit=yfit)
    residual = y – yfit
    threshold = sigma * stddev(residual) ; sigma = 2
    outliers = where(abs(residual) > threshold)
    If size(outliers)/size(x) < ratio then break ; ratio = 3%
    ; else prepare for the next iteraiton
    valid = where(abs(residual) <= threshold)
    x = x[valid]
    y = y[valid]
  endfor

## Slide 18: Recursive filtering result
Improved relative calibration accuracy
method with less variation contributes (weights) more in the RF results
 | #Obs. | #Outliers | %Selection
Desert | 85 | 37 | 56.5%
Ray-matching | 85 | 8 | 90.6%
DCC | 85 | 7 | 91.8%
Moon | 48 | 15 | 68.8%

## Slide 19: DCC + Lunar for GOES-12
DCC
  Monthly median
Lunar calibration with GIRO model
  All the observations available
 | # Obs. | # Outliers | % Selection
DCC | 85 | 7 | 91.8%
Moon | 48 | 15 | 68.8%

## Slide 20: sensitivity
Combined Vicarious Calibration

## Slide 21: Sensitivity Analysis (Desert + Ray-matching+ DCC + Moon)
Trending of the integrated method is generally the closest to that of DCC, followed by that of Ray-matching, moon and then desert method
Notes: The iteration would break after 4th loop when the number of outliers is less than 3%. When the iteration is forced to continue in the sensitivity test, Fig a shows that the fitting residuals is nearly symmetrically distributed (skewness close to 0), the standard deviation of the fitting residual (STDV) becomes stable, and fewer outliers were found after the 4th iteration (none after the 10th).
Fig b shows that the fitting functions are very similar for all the 14 loops.
Figure c: M4 stands for the recursive filtering results from the four methods (Desert + Raymatching + DCC + Moon). The iteration breaks at 4th loop as shown in Figure a.
Figure d: solid lines are the trending of each individual method, referred to the primary y-axis. The dash lines are the trending difference between each individual method to the integrated method (M4), referred to the second y-axis. Trending of the integrated method is generally the closest to that of DCC, followed by that of Ray-matching, moon and then desert method

## Slide 22: Sensitivity Analysis (DCC + Moon)
Same as previous slide. Only for DCC and Moon combined results
Sensitivity of threshold for the outlier selections in the recursive filtering was examined. The Sigma value is tested with 1.5, 2.0, 2.5, and 3. No difference in the outlier selections was found in this case.
DCC weights much more than lunar method in the integrated method result

## Slide 23: Limited Data in Operation
Fitting function: y = a + b* t + c * t^2
Fitting parameters are noisy in the first year and become stable after about two years
Using the long-term trending as the reference, the fitting error of near real-time data is in general less than 2% after first year and reduced to about 1% after 2-3 years.
Notes: Time-series of trending coefficients and the error budgets of integrated method for the ‘near real-time’ data. The coefficients are calculated with the fitting functions (Equation 2) in which the unit of time t is year after operation. The time on x-axis starts with the GOES-12 operational date on April 1, 2003. No integrated method results are available in the first three months.

## Slide 24: Absolute calibration
Combined Vicarious Calibration

## Slide 25: From Relative to Absolute Calibration
Use of desert and DCC methods as example to calculate the absolute calibration correction coefficients
Re-construct the pre-launch reflectance from the integrated calibration result.

## Slide 26: Absolute Calibration Correction
Sonoran Desert:
DCC :
The systematic difference between desert and DCC method is less than 1%

## Slide 27: conclusions
Combined Vicarious Calibration

## Slide 28: Discussions & Conclusions
Recursive filter can combine multiple calibration results
  Normalize, then identify and remove outliers recursively
DCC / Lunar calibration contributes the most / least
  As expected, considering their variations.
  Had the similar effect as intended by weighted average.
  A priori information is still useful, since recursive filter works better when population is nearly normal.
In operation, it take ~2 years for reliable results.
  Relatively stable fitting parameters.
  Uncertainty <1-2%
Recommend using DCC for the absolute calibration
  1% absolute calibration difference between DCC and desert.
  GOES views the Sonoran desert at a large angle – GOES scan mirror reflectivity may depend on incident angle.