## Slide 1: Application of GIRO to Himawari-8/AHI
Hidehiko Murata and Masaya Takahashi
Meteorological Satellite Center,
Japan Meteorological Agency
Himawari-8
Himawari-9
(will be launched in 2016)
In collaboration with EUMETSAT and NOAA

## Slide 2: Contents
Introduction
Himawari-8 and AHI
Lunar observation/calibration
Overview of the Moon observation
First results of the AHI lunar calibration
Summary and future plans

## Slide 3: Himawari-8/AHI VIS/NIR bands
AHI VIS/NIR spec.
Himawari-8/AHI
GOES-R/ABI
Meteosat-10/SEVIRI
Band | Central Wavelength [μm] | Spatial Resolution [km] | # of detectors (*) | Digital count quantization level [bit]
1 | 0.47 | 1 | 676 (3) | 11
2 | 0.51 | 1 | 676 (3) | 11
3 | 0.64 | 0.5 | 1460 (3) | 11
4 | 0.86 | 1 | 676 (6) | 11
5 | 1.6 | 2 | 372 (6) | 11
6 | 2.3 | 2 | 372 (6) | 11
0.4 0.6 0.8 1.0 1.2 1.4 1.6 1.8 2.0 2.2 2.4 [µm]

## Slide 4: Timeline of JMA GEO satellites operation

## Slide 5: Himawari-8/9 AHI observations in 10 minutes time frame
These “super-rapidscans” (every 30 seconds)
will also be used for the lunar observations
Full Disk Observation
every 10 min.
Region 1
(NE Japan)
Every 2.5 min.
Region 2
(SW Japan)
Every 2.5 min.
Region 3
(Target Area)
Every 2.5 min.
Region 4
(Landmark Area)
Every 30 sec.
Region 5
(Landmark Area)
Every 30 sec.

## Slide 6: Observation timeline (10 min.)
(based on Yokota and Sasaki 2013)
[sec]
Full Disk | Region 1 (NE Japan)
BB Obs. | Region 2 (SW Japan)
Space Look | Region 3 (Target Area)
 | Region 4 (Landmark Area)
 | Region 5 (Landmark Area)
One scan (swath)
[sec]

## Slide 7: Himawari-8/AHI lunar observation/calibration

## Slide 8: Himawari-8/AHI lunar observation
Band1 (0.47μm)
2014-12-28T19:05:24-19:56:55
Band5 (1.6μm)
2015-01-12T08:29-18-08:34:27
The Moon acquisition by super-rapidscan mode / 30 sec
Performed based on the Moon position prediction
  Dec 2014: fixed observation position within the 10 minutes obs. frame
  Jan 2015~: predict the Moon position for each observation
~400 lunar observations applicable to GIRO in Dec 2014 and Jan 2015 (have been already shared w/ EUMETSAT)
Lunar observation netCDF
Will be shared within the lunar calibration community
One swath for the Moon observation
Pre-operational phase: Dec. 2014 - mid 2015
Moon
Observation domain of super-rapidscan
Targeting observation size: ~0.8 x 1.6 [deg]
Moon apparent diameter: ~0.5 [deg]

## Slide 9: Himawari-8/AHI lunar observation
The Moon acquisition by super-rapidscan mode / 30 sec
Performed based on the Moon position prediction
  Dec 2014: fixed observation position within the 10 minutes obs. frame
  Jan 2015~: predict the Moon position for each observation
~400 lunar observations applicable to GIRO in Dec 2014 and Jan 2015 (have been already shared w/ EUMETSAT)
Lunar observation netCDF
Will be shared within the lunar calibration community
One swath for the Moon observation
Pre-operational phase: Dec. 2014 - mid 2015
Moon
Observation domain of super-rapidscan
Targeting observation size: ~0.8 x 1.6 [deg]
Moon apparent diameter: ~0.5 [deg]
2015 GSCIS annual meeting, 16-20 March, 2015, New Delhi, India
Band1 (0.47μm)
2014-12-28T19:52:24-19:56:55
Band5 (1.6μm)
2015-01-12T08:29-18-08:34:27

## Slide 10: AHI lunar observation – the Moon positions in a imagette
2014-12-29T20:44:56
Band1
(0.47um)
Band2
(0.51um)
Band3
(0.64um)
Band4
(0.86um)
Band5
(1.6um)
Band6
(2.3um)
ITT EXELIS(2013): AHI-8 EQUIPMENT MANUAL 2
CALIBRATION AND ALIGNMENT HANDBOOK Rev. J
Alignment of the AHI VIS/NIR bands
Band number

## Slide 11: Time series of AHI lunar observations
The Moon phase angles applicable to the GIRO will be covered in the next few months
Observations within one day
~ 1 month
GIRO (ROLO) applicable range

## Slide 12: Histogram of AHI digital counts
2014-12-16T09:05:56Z
Band1
(0.47um)
Band2
(0.51um)
Band3
(0.64um)
Band4
(0.86um)
Band5
(1.6um)
Band6
(2.3um)
The Moon mask threshold will be investigated further

## Slide 13: Himawari-8/AHI lunar calibration: First results
Difference of lunar irradiance between the observation and GIRO [%], rescaled at the first result
Time series of ΔIrr for Band3 (0.64um)
Δirr (Y-axis) vs. lunar phase angle (X-axis) for Band5 (1.6um)
16 days
MSG1/SEVIRI NIR1.6 result
by Bart at the Lunar calibration Workshop in Dec. 2014
NOTE: time scale is completely different between these two instruments’ results

## Slide 14: Summary and future plans
Himawari-8
Operation: mid 2015~
  All the Moon images within the FoR would be observed
AHI lunar observation netCDF: to be shared within the lunar calibration community
Future plans
Comparison with other calibration results
  On-board calibration by a solar diffuser
  DCC inter-calibration, RTM vicarious calibration, Ray-matching, etc.
Investigation of
  GIRO (ROLO) irradiance stability (relative precision)
  phase angle dependence of the GIRO (ROLO) irradiance
  scan angle dependence on the lunar calibration

## Slide 15
Thanks for your attention!

## Slide 16: JMA geostationary satellites
GMS (Geostationary Meteorological Satellite)
MTSAT (Multi-functional Transport SATellite )
Back-up operation of GMS-5 w/ GOES-9 by NOAA/NESDIS:
2003/05/22-2005/06/28
(GOES-9)
Aug 1981
GMS-2
(Himawari-2)
Aug 1984
GMS-3
(Himawari-3)
Sep 1989
GMS-4
(Himawari-4)
Mar 1995
GMS-5
(Himawari-5)
Feb 2005
MTSAT-1R
(Himawari-6)
Feb 2006
MTSAT-2
(Himawari-7)
Jul 1977
GMS
(Himawari)
Himawari
2014
Himawari-8
2016
Himawari-9
ⒸNASDA
ⒸNASDA
ⒸNASDA
ⒸNASDA
ⒸNASDA
ⒸSS/L
ⒸMELCO
Satellite | Operation period
GMS | 1978 – 1981
GMS-2 | 1981 – 1984
GMS-3 | 1984 – 1989
GMS-4 | 1989 – 1995
GMS-5 | 1995 – 2003
GOES-9 | 2003 – 2005
MTSAT-1R | 2005 – 2010
MTSAT-2 | 2010 – 2015
Himawari-8 | 2015 – 2022
Himawari-9 | 2022 – 2029
launched on 7 Oct. 2014
Operation: mid 2015~