## Slide 1
Rűdiger Lang, Rose Munro, Gabriele Poli, Antoine Lacan, Christian Retscher, Michael Grzegorski, Alexander Kokhanovsky, Rasmus Lindstrot, Richard Dyer and Ed Trollope
GSICS UV sub-group: GOME-2 hyper-spectral radiances for calibration of AVHRR imager data

## Slide 2
GOME-2 Hyper-Spectral Radiances for calibration of AVHRR imager data - Outline
GOME-2 calibration of AVHRR channel 1 radiances
GOME-2 calibration of Seviri channel 1 radiances (as additional material)
PMAp – Polar Multi-sensor Aerosol product: Towards continuous calibration of AVHRR Visible (and IR) channels

## Slide 3
Metop-A/B GOME-2 Radiometric accuracy and calibration
Earthshine measurement co-location FM2 to FM3 / Co-location criteria
Co-location criteria:
Area overlap > 80%
Geo. AMF diff < 2%
Relative O2 A-Band residual (ROR) smaller than empirical threshold
GOME-2 FM3 Metop-A
GOME-2 FM2 Metop-B
Metop-A (M02): FM3
Metop-B (M01): FM2
Time difference:
49 min (half an orbit)

## Slide 4
Metop-B / FM2 Radiometric accuracy and calibration
Earthshine measurement co-location to Metop-A / FM3 / Radiances
Radiance residuals are dominated by FM3 degradation signature
GOME-2 FM3 Metop-A
GOME-2 FM2 Metop-B

## Slide 5
Example for an average over multiple FM2/FM3 residuals in reflectance.
Prelim. Approach:
A linear background is subtracted per channel to account for the difference in reflectance-degradation and for broad-band differences in the different observed atmospheric paths per instrument.
Looking for the small scale structures per channel
GOME-2 FM3 Metop-A
GOME-2 FM2 Metop-B
Metop-A/B GOME-2 Radiometric accuracy and calibration
Earthshine co-located Metop-A/B residuals
Background predominantly from Metop-A degradation

## Slide 6
Residual in reflectance
+ Lin. background subtracted per channel
+ 20 pix. moving average smoothing
+ Relative Oxygen-A band Residual (ROR) selection criterium
Metop-A/B GOME-2 Radiometric accuracy and calibration
Earthshine co-located Metop-A/B reflectances and residuals
GOME-2 FM3 Metop-A
GOME-2 FM2 Metop-B

## Slide 7
GOME-2 FM3 Metop-A
GOME-2 FM2 Metop-B
Metop-A/B GOME-2 Radiometric accuracy and calibration
Earthshine co-located Metop-A/B residuals / Background-subtracted
plus degradation residual from Metop-A
Background subtracted co-located Metop-A/B residual
Background subtracted temporal degradation residual Metop-A

## Slide 8
GOME-2 FM3 Metop-A
GOME-2 FM2 Metop-B
Background subtracted co-located Metop-A/B residual
without the temporal degradation contribution of Metop-A
Metop-A/B GOME-2 Radiometric accuracy and calibration
Earthshine co-located Metop-A/B residuals / Background-subtracted
plus degradation residual from Metop-A subtracted

## Slide 9
GOME-2 Ch:3 (Band 5)
1024 measurements
GOME-2 Ch:4 (Band 6)
1024 measurements
AVHRR ch1
response function
GOME-2/AVHRR reflectance inter-calibration
AVHRR ch 1/ Metop-A
GOME-2 FM3 Metop-A

## Slide 10
2007
2009
AVHRR channel 1 to GOME-2/Metop-A gain in reflectance >10%
(AVHRR < GOME-2)
Results indicate that the gain has been quite stable from 2007 to 2011.
see GSICS Quarterly Newsletters, vol 5, number 3, Latter et al.
GOME-2 FM3 Metop-A
GOME-2 / AVHRR reflectance inter-calibration
AVHRR ch 1/ Metop-A
2011
spatial aliasing accounted for

## Slide 11
GOME-2 FM3 Metop-A
2013
AVHRR channel 1 to GOME-2/Metop-A gain in reflectance <8%
(AVHRR < GOME-2)
2011 - 2013
see GSICS Quarterly Newsletters, vol 5, number 3, Latter et al.
GOME-2 / AVHRR reflectance inter-calibration
AVHRR ch 1/ Metop-A
spatial aliasing accounted for

## Slide 12
GOME-2 FM3 Metop-A
GOME-2 Metop-A degradation component modelling
Sahara – Reflectance Degradation / AVHRR ch1-domain
Reflectance
Degradation
600 nm
640 nm
680 nm
Average over all
GOME-2 viewing angles
R2 campaign
Jan 2007– Jan 2012
OPS phase (same PPF 5.3)
Jan 2012 – Oct 2013

## Slide 13
GOME-2 FM3 Metop-A
GOME-2 Metop-A degradation component modelling
Sahara – Reflectance Degradation / AVHRR ch1-domain
Reflectance
Degradation
640 nm
Over all
GOME-2 viewing angles
at
2008/02/01
2009/10/01
2011/12/01
R2 campaign
Jan 2007– Jan 2012
OPS phase (same PPF 5.3)
Jan 2012 – Oct 2013

## Slide 14
GOME-2 FM3 Metop-A
AVHRR channel 1 to GOME-2/Metop-A gain in reflectance <8%
(AVHRR < GOME-2)
2013
Corrected /
Un-corrected
see GSICS Quarterly Newsletters, vol 5, number 3, Latter et al.
GOME-2/AVHRR reflectance inter-calibration
AVHRR ch 1/ Metop-A
spatial aliasing accounted for
2013

## Slide 15
GOME-2 FM3 Metop-A
AVHRR channel 1 to GOME-2/Metop-A gain in reflectance <8%
(AVHRR < GOME-2)
2013
Corrected /
Un-corrected
see GSICS Quarterly Newsletters, vol 5, number 3, Latter et al.
GOME-2/AVHRR reflectance inter-calibration
AVHRR ch 1/ Metop-A
spatial aliasing accounted for
2013

## Slide 16: PMAp: Aerosol Optical Depth and volcanic ash from GOME-2 (Polarisation Measurements) / IASI and AVHRR
GOME-2 FM3 Metop-A
GOME-2 FM2 Metop-B
AOD / volcanic ash retrieval from two Metops – PMAp
Multi-sensor aerosol properties retrieval using GOME-2 stokes fractions from PMDs
GOME-2 PMD -
ground pixel sizes
Metop-B: 10 by 40 km
Metop-A: 5 by 40 km
GOME-2 PMD
AVHRR
GOME-2 PMD
IASI

## Slide 17: PMAp – Multi-sensor co-location for aerosol optical depth retrievals
Instrument |  | Spatial resolution | Spectral range | comments
GOME | Main science channel | 80 x 40 km | 240nm -800nm, res. 0.25-0.5nm | AAI, low spatial resolution, not used
 | Polarization Monitoring Device | 10 x 40 km Metop-B 5 x 40 km Metop-A | 311nm-803nm, 15 bands | AOD, aerosol type, AAI
AVHRR | - | 1.08 x 1.08 km | 580nm-12500nm, 5 bands | Clouds, scene heterogeneity, dust/ash
IASI | - | 12km (circular) | 3700–15500nm, resolution 0.5 cm-1 | desert dust, volcanic ash aerosol heights
Auxiliary data | ECMWF wind speed (forecasting) | Temporal interpolation necessary | - | Required for retrievals over ocean
 | surface albedo, Surface elevation | - | - | Required for retrievals over land
Target spatial resolution
GOME-2 FM3 Metop-A
GOME-2 FM2 Metop-B

## Slide 18: PMAp – Multi-sensor co-location for aerosol optical depth retrievals The GOME-2 PMD / AVHRR / IASI co-location
GOME-2 FM3 Metop-A
GOME-2 FM2 Metop-B
AVHRR collocation: an AVHRR pixel is collocated to a GOME-2 PMD pixel if it is inside the GOME-2 pixel
IASI collocation: a IASI pixel may span up to 6 GOME-2 PMD pixels. Which GOME-2 pixel should it be collocated to?
Spatial aliasing due to read-out timing of detectors is taken into account!

## Slide 19: PMAp – Multi-sensor co-location for aerosol optical depth retrievals Towards continuous calibration of AVHRR channels
Channel | Central wave-length[μm] | Wavelength range [μm]
1 | 0.630 | 0.580 - 0.680
2 | 0.865 | 0.725 - 1.000
3A | 1.610 | 1.580 - 1.640
3B | 3.740 | 3.550 - 3.930
4 | 10.800 | 10.300- 11.300
5 | 12.000 | 11.500- 12.500
AVHRR channels
For PMAp a consistent and continuous AVHRR – GOME-2/IASI radiometric
inter-calibration is needed!
AVHRR vs GOME-2:
Use GOME-2
Main channels to calibrate AVHRR
ch 1 and 2 (?)
Correct PMDs using MSC
GOME-2
PMD vs Main channels (MSC):
Make PMDs consistent
with main science
channel radiances
AVHRR vs IASI:
Use IASI to calibrate AVHRR ch 3 to 5
1)
2)
3)
GOME-2 FM3 Metop-A
GOME-2 FM2 Metop-B

## Slide 20
Summary
GOME-2 inter-calibration (AVHRR/Seviri/PMAp)
GOME-2 / Metop-A channel 3/4 vs AVHRR / Metop-A channel 1
Scale factor of 0.89
(2007 to 2011, 2011 to 2013 after correction of GOME-2 data, spatial aliasing accounted for by empirical correlation optimisation)
GOME-2 / Metop-B channel 3/4 vs Seviri / MSG-3 channel 1
Scale factor of ~0.92 (see additional material)
(2013, no correction of GOME-2 data, no spatial aliasing)
  PMAp: GOME-2/(IASI) AVHRR inter-calibration
Towards continuous (operational) calibration for consistency with GOME-2 main science channels

## Slide 21
Additional Material

## Slide 22
GOME-2 Ch:3 (Band 5)
1024 measurements
GOME-2 Ch:4 (Band 6)
1024 measurements
Seviri ch1
response function
GOME-2 / Seviri reflectance inter-calibration
SEVIRI ch1/ MSG-2 – preliminary results
GOME-2 FM3 Metop-A

## Slide 23
Co-location of GOME-2 Metop-A with Seviri / MSG data.
Spatial collocation: Average of all Seviri measurements (~4km) in one GOME-2 ground pixel (40 by 80 km)
Temporal collocation: Within +- 15 min of sensing time.
No spatial aliasing correction
GOME-2 FM3 Metop-A
GOME-2 / Seviri reflectance inter-calibration
SEVIRI ch1/ MSG-2 – preliminary results

## Slide 24
GOME-2/Metop-B reflectance inter-calibration
SEVIRI ch1/ MSG-3 – preliminary results
GOME-2 FM2 Metop-B
Reflectance MSG3 / Seviri Ch 1
1.5/1.5 degrees box at 0/0o
Reflectance GOME-2 / Metop-B
1.5/1.5 degrees box at 0/0o
Residual MSG-3 /Seviri vs GOME-2 / Metop-B [%]
1.5/1.5 degrees box at 0/0o

## Slide 25
GOME-2/Seviri reflectance inter-calibration
SEVIRI ch1/ MSG-3 vs GOME-2/Metop-B – preliminary results
GOME-2 FM2 Metop-B
MSG-3 / Seviri ch 1 to GOME-2/Metop-B gain in reflectance ~8%
(Seviri < GOME-2)
1st Jan to 22nd Nov 2013
* bi-weekly average

## Slide 26
GOME-2/Metop-B reflectance inter-calibration
SEVIRI ch1/ MSG-3 – preliminary results
GOME-2 FM2 Metop-B
No spatial aliasing correction yet
No correlation with SZA found
MSG-3 / Seviri ch 1 to GOME-2/Metop-B gain in reflectance ~8.5%
(Seviri < GOME-2)
1st Jan to 22nd Nov 2013