## Slide 1
NMSC/KMA
COMS MI
Visible channel Calibration
Tae-Hyeong OH, Dohyeong KIM, Hyesook LEE and Minju GU
2016 GSICS Annual Meeting (29 Feb ~ 4 Mar 2016 )@ Tsukuba, Japan

## Slide 2: GSICS Activity of KMA for visible channel
KMA Installed vicarious calibration system for visible channel using 5 targets.
  Ocean, Desert, Water Cloud, deep convective cloud (DCC), and Moon
We have tested with these target data since 2011.
Target | Ocean | Desert | Water cloud | Deep convective cloud | Moon
Period | Sep. 2011 ~ Present |  |  |  | 
Source | Regular observation |  |  |  | Special obs. (twice a month)
Obs. | - Pacific Ocean - Indian Ocean | - Simpson Desert in Australia | - Over ocean regions - Overcast clouds | - High reaching overcast clouds | - moon
Model | 6S |  | SBDART |  | GIRO

## Slide 3: GSICS Activity of KMA for visible channel
Result from DCC target is consistent with other target data, making linear regression line with high correlation.
The degradation are about 5.80% (1.23%/year) from the moon and 5.47% (1.17%/year) from the DCC from Apr. 2011 to Dec. 2015.
Result of GSICS is displayed on KMA’s homepage
  http://nmsc.kma.go.kr/html/homepage/en/gsics/gsicsMain.do

## Slide 4: Vicarious calibration algorithm using DCC
Satellite data
Well-calibrated
IR band (10.8 μm)
TB11
Selected DCC pixels
Assumption of input parameters
COT=200, re = 20 μm,
Cloud height = 1~15km
Tropical profiles
Simulated visible radiances
Yes
No
Observed visible radiances
Vicarious calibration
Cloud Radiative transfer modeling
Visible band (0.67 μm)
Same FOV
Threshold conditions
1
2
3
Sohn et al. 2009, Ham and Sohn 2010

## Slide 5: Part 1: Selection of DCC targets
Threshold conditions to select the overshooting DCCs
  convective clouds whose tops extended from 14 to 19 km with extremely high reflectivity
Solar geometry |  | SZA ≤ 40º
Viewing geometry |  | VZA ≤ 40º
 | - geometry criteria: to minimize navigation errors and 3-D radiative effects | 
Surface type |  | No restriction
 | - no restriction of surface type: less influence on reflectance of DCC targets | 
Cloud conditions (1) |  | target pixel: TB11 ≤ 190 K
 | - only temperature criterion: overshooting DCCs represent cloud top temperature lower than the TTL temperature (~190 K) | 
Cloud conditions (2) |  | environmental pixel: STD(TB11) ≤ 1K
 |  | environmental pixel: STD(R0.6)/Mean(R0.6) ≤0.03
 | - two types of homogeneity checks: to avoid selection of cloud edge or small-scale plumes | 
target pixel
9 x 9
environmental pixels
STD,
STD/Mean
TB11: brightness temperature at 10.8 μm, R0.6: reflectance at 0.67 μm, STD: standard deviation in environmental (9x9) pixels

## Slide 6: Part 2: cloud RTM description (1/2)
SBDART (Santa Barbara Disort Radiative Transfer) model (Ricchiazzi et al. 1998) - based on DISORT (Discrete Ordinates Radiative Transfer) model - capable up to 32 streams - relatively accurate and efficient RTM for cloudy atmosphere
KMA used some options as follows:
  phase function: delta-fit method (Hu et al. 2000) - bulk phase function: strong forward peak and thousands of Legendre polynomials - needed to truncate method for the phase function in the forward direction - to reduce the computational burden without degrading accuracy
  gases absorption: correlated-k-distribution (CKD) method (Kratz 1995; Kratz and Rose 1999) - gaseous absorption associated with Rayleigh scattering
  sfc. info.: oceanic surface properties for any surface type

## Slide 7: Part 2: cloud RTM description (2/2)
Scattering properties: Baum model Baum et al. (2005a and 2005b) for non-spherical ice particles
  - based on in-situ measurements to obtain habit fractions (Heymsfield et al., 2002) - use single scattering properties of droxtals, hexagonal plates, hollow columns, solid columns, bullet rosettes, and aggregates
  (Yang and Liou, 1996a and 1996b; Yang et al., 2003a and 2003b) - band averaged scattering properties with respect to re by integration of single scattering properties
    Qext (extinction efficiency), ω0 (Single scattering albedo),
    g (asymmetry factor), P(Θ) (phase function), fd (delta transmitted energy)

## Slide 8: Part 2: RTM inputs
# of streams | 20 | Cloud conditions | Ice phase (use Baum scattering model)
Simulation wavelength | 0.57488 μm – 0.77992 μm |  | COT = 200
 | (interval: 0.04101μm) |  | Effective radius = 20 μm
Filter function | rectangular filter |  | Cloud top-height = 15 km
Geometry | SZA, VZA, SAA, VAA (Observed pixels information) |  | Cloud depth = 14 km
surface type (BRDF) | Ocean |  | 
atmosphere | Tropical standard profile |  | 
Input |  | Sfc. | Atmos. | AOT | Cloud |  |  | 
Parameters |  | albedo | profiles | At 0.55 μm | COT | re | Zc | Zc
Reference value |  | Oceanic | TRO | 0 | 200 | 20 μm | 15 km | 14 km
Input range |  | 0 – 0.4 | MLS | 0 – 3 | 100 – 400 | 10 – 30 | 12 – 18 | 10 – 14
Maximum uncertainty | SZA = 0º | 0.09% | 1.29% | 0.14% | 4.79% | 1.64% | 0.15% | 0.24%
 | SZA = 10º | 0.09% | 1.29% | 0.14% | 4.79% | 1.96% | 0.16% | 0.24%
 | SZA = 20º | 0.09% | 1.29% | 0.14% | 4.74% | 2.35% | 0.18% | 0.24%
 | SZA = 30º | 0.09% | 1.30% | 0.14% | 4.60% | 3.02% | 0.21% | 0.24%
 | SZA = 40º | 0.08% | 1.30% | 0.13% | 4.41% | 3.02% | 0.21% | 0.24%
uncertainty ranges
in Appendix of Sohn et al. (2009)
RTM input parameters for COMS calibration using DCC target

## Slide 9: Results
SBDART model could be constructed bi-directional reflectance distribution structure above cloud layer with respect to various input parameters and wavelength of channels, not implying Lambertian surface.
DCC BRDF is highly dependent on scattering parameterization of RT model.
Vicarious calibration for COMS using DCC target has been conducted well in spite of dependency on the result of RT model simulation.
Result from DCC target is consistent with other target data and the degradation are about 5.47% (1.17%/year) from Sep. 2011 to Dec. 2015.

## Slide 10: NASA DCC method for COMS MI
DCC selection thresholds : same as Doelling et al. (2013) except for homogeneity check for MODIS data
Use nominal date (MYD021KM) in this analysis
DCC domain: (COMS : 128.2〫E)
  20〫S < latitude < 20〫N
  108.2〫E < longitude < 148.2〫
COMS DCC threshold | parameter | NASA DCC threshold
TB11 ≤ 190 K | Window brightness temperature | BT11μm < 205°K
STD(TB11) ≤ 1K | Brightness temperature homogeneity | Standard deviation of 3x3 pixels BT11μm < 1°K
STD(R0.6)/Mean(R0.6) ≤0.03 | Visible radiance homogeneity | Standard deviation of 3x3 pixels visible radiance < 3%
SZA < 40° | Solar zenith angle | SZA < 40°
VZA < 40° | View zenith angle | VZA < 40°
 | Local time range at GEOSat longitude | 12:00 PM < image time < 3:00 PM

## Slide 11: Comparison other results
Comparison with Lunar, DCC(RTM) and DCC(NASA method)
  Period : April 2011 ~ December 2015
  DCC(RTM) : 1.1689/year, Moon : 1.2318/year, DCC(NASA):0.9757/year
  Need more investigation for the DCC result with NASA method

## Slide 12: Summary and Future Plans
Summary
Implementation of NASA’s DCC method to COMS MI
- good agreement with the ATBD
Need more investigation to interpret the large standard deviation
Future plans
Investigate the impact of BRDF and SBAF on the calibration results
More test to compare the results using Sohn’s and NASA method for evaluating the uncertainty

## Slide 13
Thank you…

## Slide 14: vicarious calibration algorithm using DCC
Sohn et al. 2009, Ham and Sohn 2010
MODIS τc,0.646
Ice
Baum Scattering Data
(Qext, g, ω, P(Θ), fd )
Water
Mie Scattering Data
(Qext, g, ω, P(Θ), fd )
MODIS
Geometry, Ts
Qext, g, ω, P(Θ), fd
Zc
MODIS Lat./Lon.
Interpolation
with respect to re
As
P(z)
MODIS Pc
RTM
Simulated Radiances
MODIS QA 1km
Phase Flag
Collocation
P(z)
T(z)
ρH2O (z)
ρO3(z)
CKD
τg
Observed Radiances
MODIS Land Cover Type
(0.05° grid data)
AIRS
MODIS Cloud Pixels
(N=1, τc,0.646≥ 10)
τc,λ
Qext
Ocean pixels

## Slide 15: Results of NASA DCC method
Good agreements with the ATBD or those of other GEO satellites
Time series of mode and mean of DCC pixels
Mode : DN = 0.02 x Day + 623.020
Mean : DN = 0.03 x Day + 582.859
Mode : DN = 0.00 x Day + 459.336
Mean : DN = 0.00 x Day + 451.500

## Slide 16: Monthly Gain
SBAF for COMS : 0.9427 (DCC web meeting in June 5, 2015)
Need to compare with the results based on RT simulation method
Usage of other SBAF or that calculated ourselves will also be considered
Gain : 2.6797e-05 x Day + 0.712