## Slide 1: Diurnal and Seasonal variations of COMS IR inter-calibration
Dohyeong Kim, and Minju Gu

## Slide 2: Diurnal Variations (COMS IR)

## Slide 3: Spatial and temperature distributions of collocation
Example of spatial and temperature distributions of COMS vs. LEO collocations on June/10/2015~June/19/2015, for COMS IR2 vs. IASI

## Slide 4: TB bias(GEO-LEO) using IASI-A,-B, AIRS & CrIS
Check the TB bias from COMS/MI and LEOs(MetOp-A,B/IASI, Aqua/AIRS and SNPP/CrIS) for four IR Channels.
 |  | IASI-A | IASI-B | AIRS | CrIS
IR1 | Number | 137840 | 109334 | 566782 | 241119
 | Bias | 0.1072 | 0.1260 | 0.0437 | 0.0266
 | RMSE | 0.3517 | 0.3489 | 0.4277 | 0.3981
 | Slope | 0.9938 | 0.9967 | 1.0032 | 1.0004
 | Intercept | 1.9238 | 1.0903 | -0.8982 | -0.0973
IR2 | Number | 151320 | 121346 | 612961 | 263226
 | Bias | -0.0073 | 0.0224 | 0.0233 | -0.0008
 | RMSE | 0.3352 | 0.3364 | 0.3346 | 0.2887
 | Slope | 0.9931 | 0.9931 | 0.9999 | 0.9987
 | Intercept | 1.9972 | 2.0201 | 0.0518 | 0.3900
IR3 (WV) | Number | 178297 | 140999 | 741373 | 325975
 | Bias | -0.8373 | -0.8544 | -1.0767 | -0.9880
 | RMSE | 0.4211 | 0.4170 | 0.4705 | 0.4369
 | Slope | 0.9828 | 0.9851 | 0.9859 | 0.9886
 | Intercept | 3.3749 | 2.7865 | 2.3656 | 1.7984
IR4 (SWIR) | Number | 43437 | 35031 | 112060 | 36243
 | Bias | 0.1508 | 0.1421 | -0.1746 | -1.4627
 | RMSE | 0.2651 | 0.2381 | 0.3380 | 0.8662
 | Slope | 0.9959 | 0.9953 | 0.9885 | 0.9450
 | Intercept | 1.3592 | 1.5193 | 3.1955 | 14.8188

## Slide 5: Trend of TB bias
The TB difference between MI and LEOs(IASI-A, -B, AIRS & CrIS) as a function of the MI TB
IR1
WV
IR2
SWIR

## Slide 6: Diurnal variation for Four Seasons (TB biases)
Diurnal variation for TB bias of IASI-A, -B, AIRS and CrIS, and for Frequency of MBCC application
IR1
WV
IR2
SWIR

## Slide 7: Diurnal Tb variation vs. MBCC performance
COMS-AIRS Diurnal Tb variation |  |  |  |  | 
 | Winter | Spring | Summer | Autumn | Total
IR1 | 0.9445 | 0.7586 | 0.7103 | 0.6082 | 0.7497
IR2 | 0.4856 | 0.4537 | 0.4817 | 0.2943 | 0.3808
WV | 0.6681 | 0.5893 | 0.5817 | 0.5677 | 0.5595
COMS-AIRS MBCC performance |  |  |  |  | 
 | Winter | Spring | Summer | Autumn | Total
IR1 | -0.3775 | -0.2784 | -0.5283 | -0.0873 | -0.3133
IR2 | -0.0130 | -0.0238 | -0.1772 | 0.1450 | -0.0142
WV | -0.4844 | -0.3706 | -0.4706 | -0.2675 | -0.3841
period: AIRS(April 2011 - Dec. 2015)
, LST
, LST
COMS-CrIS Diurnal Tb variation |  |  |  |  | 
 | Winter | Spring | Summer | Autumn | Total
IR1 | 0.8201 | 0.6734 | 0.7468 | 0.7063 | 0.6640
IR2 | 0.3875 | 0.2467 | 0.3670 | 0.2835 | 0.2456
WV | 0.7099 | 0.5294 | 0.5494 | 0.5519 | 0.5395
COMS-CrIS MBCC performance |  |  |  |  | 
 | Winter | Spring | Summer | Autumn | Total
IR1 | -0.3907 | -0.2711 | -0.5455 | -0.1575 | -0.3321
IR2 | -0.0611 | 0.0166 | -0.1601 | 0.1079 | -0.0179
WV | -0.4621 | -0.2801 | -0.4723 | -0.2655 | -0.3662
*Reference: Fangfang Yu et. al(2013), Diurnal and Scan Angle Variations in the Calibration of GOES Imager Infrared Channels

## Slide 8: Slope retrievals (morg vs. mest(MBCC) vs. mGSICS)
()
: true reference
: true value
(regression coeff.)
,
, : original calibration coeff.
Original slope :
GSICS Corrected slope(?)
Estimated slope (MBCC) :
,
where : “predicting” temperature
(primary mirror temperature)
mGSICS
mest(MBCC)
morg
Collaboration with NOAA (Fred and Fangfang)

## Slide 9: Slope retrievals (mGSICS)
  [Wm-2str-1m-1], [mWm-2str-1cm-1]
  To retrieval
  [Wm-2str-1m-1]
(: given value)
, )
 =
[Wm-2str-1m-1]
Least square fitting method is applied to every scene
If there are more than 50 of matching, and are obtained through linear fit
The following figure is an example of scene for 16: 30 on March/13/2016
Linear Regression (red line)
Slope:
Intercept :

## Slide 10: Slope Comparison (morg vs. mest(MBCC) vs. mGSICS)
Period : March/12/2016~April/10/2016, June/07/2016/~July/05/2016, Sep./10/2016~Oct./11/2016, Dec./10/2016~Jan./11/2017
under thermal stress(22:00~03:00 LST) [Wm-2str-1m-1/count]
mGSICS
mest(MBCC)
morg
IR1
WV
IR2
SWIR

## Slide 11: Time series for Slope
IR1 | March/12 -April/10 | June/7 -July/5 | Sep/10 -Oct11 | Dec/10 -Jan/11
Ori.mean | -1.55029e-2 | -1.58900e-2 | -1.55087e-2 | -1.54963e-2
Ori.stdv | 0.0905e-3 | 0.1814e-3 | 0.1977e-3 | 0.1134e-3
Est.mean | -1.55551e-2 | -1.59138e-2 | -1.55311e-2 | -1.54904e-2
Est.stdv | 0.1305e-3 | 0.1574e-3 | 0.204e-3 | 0.1079e-3
GSICS mean | -1.88392e-2 | -1.89185e-2 | -1.87313e-2 | -1.87766e-2
GSICS stdv | 0.3723e-3 | 0.5406e-3 | 0.4871e-3 | 0.4167e-3
IR2 | March/12 -April/10 | June/7 -July/5 | Sep/10 -Oct11 | Dec/10 -Jan/11
Ori.mean | -1.32695e-2 | -1.36696e-2 | -1.32526e-2 | -0.29184e-2
Ori.stdv | 0.0761e-3 | 0.1150e-3 | 0.1020e-3 | 0.0081e-3
Est.mean | -1.33140e-2 | -1.37036e-2 | -1.32597e-2 | -0.29327e-2
Est.stdv | 0.0819e-3 | 0.0827e-3 | 0.0991e-3 | 0.0059e-3
GSICS mean | -1.66532e-2 | -1.66500e-2 | -1.64663e-2 | -0.35651e-2
GSICS stdv | 0.5170e-3 | 0.5072e-3 | 0.3240e-3 | 0.2424e-3
mGSICS
mest(MBCC)
morg
Mar-Apr
Jun-Jul
Sep-Oct
Dec-Jan

## Slide 12: Time series for Slope
WV | March/12 -April/10 | June/7 -July/5 | Sep/10 -Oct11 | Dec/10 -Jan/11
Ori.mean | -1.62584e-2 | -1.65747e-2 | -1.62387e-2 | -1.62236e-2
Ori.stdv | 0.1056e-3 | 0.2154e-3 | 0.2196e-3 | 0.1430e-3
Est.mean | -1.62900e-2 | -1.66039e-2 | -1.62412e-2 | -1.61902e-2
Est.stdv | 0.1383e-3 | 0.1851e-3 | 0.2187e-3 | 0.1098e-3
GSICS mean | -1.81364e-2 | -1.80508e-2 | -1.80778e-2 | -1.80601e-2
GSICS stdv | 0.5543e-3 | 0.5956e-3 | 0.4137e-3 | 0.3193e-3
SWIR | March/12 -April/10 | June/7 -July/5 | Sep/10 -Oct11 | Dec/10 -Jan/11
Ori.mean | -0.28924e-2 | -0.28331e-2 | -0.29511e-2 | -0.29184e-2
Ori.stdv | 0.0528e-3 | 0.0082e-3 | 0.0322e-3 | 0.0081e-3
Est.mean | -0.29240e-2 | -0.29132e-2 | -0.29604e-2 | -0.29327e-2
Est.stdv | 0.0075e-3 | 0.0075e-3 | 0.0136e-3 | 0.0059e-3
GSICS mean | -0.34766e-2 | -0.34605e-2 | -0.34613e-2 | -0.35651e-2
GSICS stdv | 0.1510e-3 | 0.1380e-3 | 0.1712e-3 | 0.2424e-3
Mar-Apr
Jun-Jul
Sep-Oct
Dec-Jan

## Slide 13
Erroneous Slope (morg – mest(MBCC))
Summer
Spring
Autumn
Winter

## Slide 14
Erroneous Slope (Autumn and Winter)
Autumn
Winter

## Slide 15: Seasonal Variations (COMS IR)

## Slide 16: Time series (TB biases and RMSE)
IR1
IR2
WV
SWIR

## Slide 17: Seasonal variation (AIRS and CrIS)
IR1
WV
IR2
SWIR

## Slide 18
Seasonal variation (IASI)
IR1
WV
IR2
SWIR

## Slide 19: Seasonal variation (AIRS/CrIS and IASI)
IR1
WV
IR2
SWIR

## Slide 20: Telescope Primary Temperature(day+night)
IASI and AIRS collocation time (daytime vs. nighttime) during 2016

## Slide 21: Telescope Primary Temperature(night)
For IR1, IR2 and WV, IASI and AIRS collocation time (nighttime, MBCC) during 2016
IASI
AIRS

## Slide 22: Scan Mirror Temperature(night)
  1 : This correction algorithm is especially pertinent for the shortwave IR channels around satellite midnight and during eclipse seasons.(GOES-8 imager midnight effects and slope correction, Weinreb)
  2 : While it is usually applied with high frequency for about 8h around satellite midnight for the short-wave channels, it may only be intensively used right after satellite midnight or even barely used for the other IR channels.(Diurnal and Scan Angle Variations in the Calibration of GOES Imager Infrared Channels, Yu)
For SWIR, IASI and AIRS collocation time (nighttime, MBCC) during 2016
IASI
AIRS

## Slide 23
Erroneous Slope (Spring and Summer)
Summer
Spring

## Slide 24
Erroneous Slope (Autumn and Winter)
Autumn
Winter

## Slide 25: Optics temperature (diurnal variation)
Spring
Summer
Autumn
Winter

## Slide 26: Summary and discussion
Diurnal and seasonal variations are not ignorable in 3 axis stabilized satellite
MBCC corrections (COMS) showed seasonal characteristics
  Spring/Autumn shows strong correction while summer/winter weak correction
Requirements from SST community
  To provide the each band(IR1), interband(IR1-IR2), and their combination correction coefficients
  To provide the diurnal variation of GSICS correction
  Need to update the corrected TB not retrieval coefficients

## Slide 27: GEO-GEO intercomparison (COMS vs MTSAT-2 COMS vs Himawari-8)

## Slide 28: COMS vs MTSAT-2
Collocation
1) Time
- 8 times/day(COMS: 15min, MTSAT-2: 33min)
* SWIR: night time(11~20UTC)
2) Optical path
- Δsec(satellite zenith angle) < 1%
- lon,lat≤each satellite center lonlat±35º
→△lon: 135.21~137.97º, △lat: -35~35º
Uniformity Test
- 9x9(LE) pixels : TB STDV(LE) < 1K
IR1 TB[K] difference in collocaion region
(1 Jun. 2014 02:15/02:30UTC)

## Slide 29: COMS vs MTSAT-2
COMS SRF(solid), MTSAT-2 SRF(dashed)

## Slide 30: COMS vs MTSAT-2
11:15UTC(COMS) 11:33UTC(MTSAT-2)
‘a’ with SBAF, ‘b’ without SBAF application
* SBAF is important
  * SBAFs obtained
  via http://cloudsgate2.larc.nasa.gov/

## Slide 31: COMS vs MTSAT-2
TB bias over TB(GSICS Corrected)
- SBAFs applied
- 3 months(Jun., Sep. Dec.) mean
- part of shading: Uncertainty(k=1)

## Slide 32: COMS vs Himawari-8
COMS Lv1b
HIMAWARI-8 Lv1b
Uniformity test
To minimize the difference of observation conditions, only use the pixel passed the Uniformity test
: when 50 of 81(9x9(LE))pixels exist, STDV(LE) < 1K
Collocation
1. Time: 8 times/day (COMS 15min, H-8 20min)
* SWIR: only night time(11~20UTC)
2. Optical path: Δsec(satellite zenith angle) < 1%, lon,lat≤±35º
- △Lat: -35~35º
- △Lon: 132.5~136.45º(about △3.94)
3. Channel(central wavelength) match
1) COMS IR1(10.79㎛) - H8 Ch13(10.41㎛)
2) COMS IR2(12.06㎛) - H8 Ch15(12.38㎛)
3) COMS WV(6.74㎛) - H8 Ch08(6.24㎛)
4) COMS IR1(3.75㎛) - H8 Ch07(3.89㎛)
COMS IR1 TB[K] in collocaion region
(1 Feb. 2017 08:15UTC)
Methods

## Slide 33: COMS vs Himawari-8
COMS SRF(solid), HIMAWARI-8 SRF(dashed)

## Slide 34: COMS vs Himawari-8
11:15UTC(COMS) 11:20UTC(H-8)
‘a’ with SBAF, ‘b’ without SBAF application

## Slide 35: Diurnal variation (with MTSAT-2)
Case 1
Case 2
Case 3
To select clear pixel during the day (00-23UTC)
IR1(10.8um), IR1-IR2(12um)
Case 1: 31 Jan. 2015 / 35.09ºN, 129.32ºE
Case 2: 31 Jan. 2015 / 39.02ºN, 122.0ºE
Case 3: 1 Feb. 2015 / 42.99ºN, 136.03ºE
COMS IR1 image (16:00UTC 31 Jan. 2015)
Notes: Case1: 2015년 1월 31일 35.09ºN, 129.32ºE COMS(992,384), MTSAT-2(1035,499)
Case2: 2015년 1월 31일 39.02ºN, 122.0ºE COMS(1118, 238), MT-2(1200, 348)
Case3: 2015년 2월 1일 42.99ºN, 136.03ºE COMS(840,308), MTSAT-2(915,428)

## Slide 36: Diurnal variation (with MTSAT-2)
IR1(10.8μm) TB bias
Apply SBAF
Bias < 1K
Warm bais is shown except for midnight
Around midnight, cold bias increase due to Midnight effect
MBCC(Midnight blackbody calibration correction) is implemented for four hours around midnight
Application of MBCC
Red: 1x1 pixel
Blue: 2x2 pixel
Black: 3x3 pixel
IR1(10.8μm) TB bias
Red: 1x1 pixel
Blue: 2x2 pixel
Black: 3x3 pixel
[Case 1] 31 Jan. 2015/35.09ºN,129.32ºE
[Case 2] 31 Jan. 2015/39.02ºN,122.0ºE
[Case 3] 1 Feb. 2015/42.99ºN, 136.03ºE
IR1-IR2 TB bias
[Case 1] 31 Jan. 2015/35.09ºN,129.32ºE
[Case 2] 31 Jan. 2015/39.02ºN,122.0ºE
[Case 3] 1 Feb. 2015/42.99ºN, 136.03ºE

## Slide 37
Thank you