## Slide 1
Spectral Correction for Inter-comparion between VIIRS and MODIS TEB
Aisheng Wu, Yonghong Li, Truman Wilson and Xiaoxiong Xiong
MODIS Characterization Support Team (MCST)
Science Systems and Application, Inc (SSAI), Lanham, MD 20706, USA
2017 GSICS GRWG/GDWG Annual Meeting, Madison, WI, Mar 20-24

## Slide 2
Early results (also presented at last GSICS) showed that real-time hyperspectral IR measurements (IASI, AIRS and CrIS) can facilitate high quality VIIRS-to-MODIS calibration and inter-comparison.
A static spectral correction reduces the BT differences from more than 2.0 K to well within 0.2 K for the spectrally matched bands at 4.0, 10 and 12µm.
One additional spectrally matched band pair at 8.5 µm was not compared due to the AIRS/CrIS spectral gap.
The current update will be focusing on the 8.5 µm bands, which corresponding to VIIRS M14 and MODIS B29.
Background
2

## Slide 3
MODIS RSR vs CrIS and AIRS spectral range
AIRS/CrIS spectral gap
3

## Slide 4
Simulated MODIS and VIIRS radiance from hyperspectral data
Spectral and spatial integration
[λ1,λ2]: range of Relative Spectral Response (RSR) wavelength
Aggregate MODIS pixels within each IASI IFOV
  MODIS 1-km pixels are collocated with an AIRS/CrIS 14-km IFOV
  Typically >100 MODIS pixels are collocated within one hyperspectral IFOV
  Only IFOVs with >70 MODIS pixels are considered in SNO data collection
MODIS Pixel
AIRS/CrIs Pixel
4

## Slide 5
VIIRS and MODIS Spectral band adjustment factor (SBAF)
SBAF=
AIRS
CrIS
  A static one-time SBAF is derived using real-time matching AIRS and CrIS hyperspectral data and the correction is radiance level dependent.
  No noticeable differences using either AIRS or CrIS in calculation of SBAF
5

## Slide 6
VIIRS and Aqua MODIS BT trends at 8.5 µm bands
  Results of both bands are referenced to a stable band at 11.0 µm
North Atlantic Ocean
6

## Slide 7
Terra MODIS BT trend at 8.5µm bands
  Results are referenced to a stable band at 11.0 µm
  The upward drift is due to steadily increasing contamination from electronic crosstalk and the derived correction is based on the lunar observations, developed by MCST. The upcoming Terra C6.1 L1B product will have the correction in place over the entire mission.
North Atlantic Ocean
7

## Slide 8
Summary
A static spectral correction currently works well between NPP VIIRS and Aqua MODIS TEB since both sensors are stably maintained.
For the spectrally matched bands at 8.5 µm between NPP VIIRS and Aqua MODIS, the spectral correction will rely on IASI data and increasing contamination in this band is possible based on the Terra performance.
Real-time hyperspectral IR measurements (AIRS, CrIS and IASI) can facilitate high quality sensor-to-sensor calibration and inter-comparison.
8