## Slide 1: Characterizing DCC as invariant calibration target
David Doelling, Rajendra Bhatt, Ben Scarino, Conor Haney, Arun Gopalan, Jack Xiong
GSICS annual meeting, Madison, WI, USA, March 20-24, 2017

## Slide 2: DCC Invariant target methodology (relative or stability calibration)
Identify monthly all DCC pixels over the domain
  The GEO and MODIS window channel IR temperatures are stable
Convert the DCC radiance to an overhead sun radiance using the Hu BRDF model
Apply an spectral band adjustment factor to the Aqua-MODIS sensor radiance to convert the radiance to an equivalent GEO sensor radiance using SCIAMACHY hyper-spectral radiances.
  very small factor for DCC <1µm
Histogram all of the pixel level DCC overhead sun radiances and determine the PDF mode radiance.
Compute the GEO calibration coefficients by monitoring the drift of the monthly GEO PDF mode radiances, which represents the visible degradation of the sensor

## Slide 3: DCC Invariant target methodology (absolute calibration)
Assume that each GEO domain has an invariant reflectance, which is dependent on the local time
  Assume that the GEO (monitored) sensor and Aqua-MODIS (reference calibration) sensor capture the same DCC over the GEO domain at the time of the Aqua-MODIS overpass
  Each domain has a distinct diurnal and seasonal cycle
  Assume the inter-annual variability of the diurnal and seasonal cycle is very small
This method makes it possible to calibrate historical as well as future GEOs
  Does not need any contemporary reference Aqua-MODIS or NPP-VIIRS radiance observations
  Able to tie historical, current and future reflected solar band calibration references

## Slide 4: Transferring the reference calibration using DCC targets
Satellite | Position | B1 | I1 | M5
Global |  | 0.4 | 0.5 | 0.4
GOES-W | 135°W | 1.0 | 0.9 | 1.0
GOES-E | 75°W | 0.5 | 0.4 | 0.5
Met-10 | 0°E | 0.7 | 0.7 | 0.7
Met-7 | 60°E | 0.9 | 1.0 | 1.1
FY2E | 86° | 0.8 | 0.7 | 0.7
MTSAT-2 | 140° | 0.8 | 0.7 | 0.7
GEO domain standard deviations
DCC mode radiances for MODIS B1, VIIRS I1 and M5 for global and GOES-W domain
• Find the DCC mode radiance reference based on Aqua-MODIS (C6) B1, NPP-VIIRS (V001) I1 and M5
• Compare the GEO domain mode radiance with respect to the global mode radiance
• Compare the reference ratio over each of the GEO domains

## Slide 5: TWP DCC inter-annual variability
270
250
230
210
CERES LW flux (Wm-2)
2000
2017
• The TWP DCC mode reflectance inter-annual (ENSO) stability is within 0.8% over the TWP

## Slide 6
Domain |  | B1 | I1 | M5 | 
 |  | GEO/global (%) | GEO/global (%) | GEO/global (%) | range
Global |  | 464.6 | 464.1 | 471.7 | 
GOES-W | 135°W | +0.5 | +0.1 | +0.0 | 0.5
GOES-E | 75°W | +0.3 | +0.2 | +0.3 | 0.1
Met-10 | 0°E | +0.5 | +0.6 | +0.6 | 0.1
Met-7 | 60°E | +0.1 | -0.2 | -0.1 | 0.3
FY2E | 86° | -0.6 | +0.3 | +0.4 | 1.0
MTSAT-2 | 140° | -0.5 | -0.5 | -0.4 | 0.1
Compare the GEO domain with the global DCC mode radiance
• Each GEO domain has a unique invariant DCC mode radiance
• The GOES-W, GOES-E, Met-10 locations have greater mode radiances
• The MTSAT-2 domain has a smaller mode radiance and has the greatest DCC frequency
• GOES-E, Met-10 and MTSAT-2 domain mode radiance is very well tied to the global

## Slide 7: Compare the reference radiances in each GEO domain
Satellite | Longitude Position | I1-B1 | M5-B1 | M5-I1
Global |  | -0.1 | 1.5 | 1.6
GOES-W | 135°W | -0.5 (-0.4) | 1.0 (-0.5) | 1.5 (-0.1)
GOES-E | 75°W | -0.3 (-0.2) | 1.4 (-0.1) | 1.7 (+0.1)
Met-10 | 0°E | -0.1 (0.0) | 1.6 (+0.1) | 1.7 (+0.1)
Met-7 | 60°E | -0.3 (-0.2) | 1.3 (-0.2) | 1.6 (+0.0)
FY2E | 86° | +0.8 (+0.9) | 2.5 (+1.0) | 1.7 (+0.1)
MTSAT-2 | 140° | -0.1 (0.0) | 1.6 (+0.1) | 1.7 (+0.1)
() Relative ratio with respect to the global
• The GEO domain reference band ratios are similar to the global for most domains except for GOES-W, and FY-2E, but mostly within the GEO domain DCC mode sigma
• Next compare the DCC mode radiance to the ray-matched calibration

## Slide 8: All-sky Tropical Ocean Ray-match calibration
Ed 3
Ed 4
GOES-13/Aqua-MODIS for April 2011
Red line = linear regression through the space clamp offset (force fit)
Black line = linear regression
Under perfect ray-matching conditions the force fit and the linear regression should be equal
• The lax angular matching, not accounting for spectral band differences, biased Ed3
Doelling et al. 2016
Decrease angle tolerance for clear-sky radiances

## Slide 9: DCC ray-matched calibration
All-sky Tropical Ocean
MTSAT-2/Aqua-MODIS Jan 2013
Doelling et al. 2016
• Find 30-km diameter cores in MODIS image with the T<220K
• Match coincident MODIS and GEO DCC pairs within a scattering angle of 15°
• Apply SBAF, and follow ATO algorithm
DCC are more Lambertian and a smaller SBAF correction and uncertainty than ATO

## Slide 10: Meteosat-10 calibration gains
• All calibration methods give consistent gains
desert

## Slide 11: Comparison of GEO calibration gains referenced to Aqua-MODIS
Satellite | Longitude | DCC-ATO (ray) | DCCmode- ATOray | DCCmode-DCCray
GOES-15 | 135°W | -0.02 | -0.37 | -0.35
GOES-13 | 75°W | -0.11 | +0.53 | +0.64
Met-10 | 0°E | 0.11 | -0.1 | -0.12
Met-7 | 60°E | 0.31 | 0.33 | +0.02
Him-8 | 86° | -0.13 | 0.47 | +0.60
MTSAT-2 | 140° | 1.02 | 0.52 | -0.51
• The DCC mode calibration is within 0.7% of either ATO and DCC ray-matching, again within the GEO domain DCC mode radiance standard deviation
• The DCC and ATO ray-matching is mostly within 0.3%

## Slide 12: Comparison of ATO and DCC ray-matched gains for Himawari-8
Him-8 band | VIIRS band | µm | Gain DCC | Stderr (%) | Gain ATO | Stderr (%) | DCC-ATO (%)
Ch1 | M3 | 0.47 | 0.3935 | 0.21 | 0.3911 | 0.31 | 0.61
Ch2 | M4 | 0.51 | 0.3676 | 0.26 | 0.3659 | 0.29 | 0.46
Ch3 | I1 | 0.64 | 0.3021 | 0.27 | 0.3013 | 0.31 | 0.27
Ch3 | M5 | 0.64 | 0.3060 | 0.25 | 0.3047 | 0.28 | 0.43
Ch4 | M7 | 0.86 | 0.1815 | 0.21 | 0.1810 | 0.27 | 0.28
Ch5 | I3 | 1.6 | 0.04420 | 0.53 | 0.04404 | 0.50 | 0.36
Ch5 | M10 | 1.6 | 0.04340 | 0.54 | 0.4331 | 0.46 | 0.21
Ch6 | M11 | 2.3 | 0.01460 | 0.39 | 0.1488 | 0.39 | -1.88
• Most Himawari-8 bands ATO and DCC ray-matched gain differences are within 0.6%

## Slide 13: Compare Aqua-MODIS B1 and NPP-VIIRS I1 over the tropics using ATO ray-matching
AZ<90°
Gain=0.9932
AZ>90°
Ocean only
AZ>90°
Gain=0.9986
All AZ
Gain=0.9951
• The global DCC mode B1-I1 is -0.1%, while the global ATO ray-matched is 0.5%
• We know that MODIS has scan angle biases, separate by azimuthal angle
• The azimuthal scan angle dependency between left and right of nadir is 0.5%

## Slide 14: MODIS scan angle dependence ray-matching impact
Pair | Ray-match | All AZ | AZ<90 | AZ>90 | AZ ratio
B1/I1 | ATO | 0.9954 | 0.9926 | 0.9993 | 0.67
 | DCC | 0.9917 | 0.9867 | 0.9959 | 0.92
B1/M5 | ATO | 0.9795 | 0.9773 | 0.9828 | 0.56
 | DCC | 0.9788 | 0.9746 | 0.9822 | 0.77
Most stderrs <0.5%
• The azimuthal scan angle dependency between left and right of nadir is 0.5 to 1.0% based on ray-matching

## Slide 15: MODIS B1 scan angle dependencies over Libya-4
Bhatt et al. 2017
• The azimuthal scan angle dependency between left and right of nadir is 0.5 to 1.0% based on ray-matching
• This is confirmed using the Libya-4 invariant target, C6 • RVS corrections were based on the DCC mode method

## Slide 16: Conclusions
The space agencies intend to maintain operational satellite VIS/NIR imagers in the 1:30 PM and 9:30 AM orbits for many decades
  The satellites will be placed 45 minutes apart in the same orbit allowing for no SNO opportunities
  Possible to inter-calibrate concurrent 1:30 and 9:30 orbits over the poles
DCC invariant target calibration can be used to transfer the calibration of one satellite with another and they need not be concurrent
  Each GEO domain DCC mode radiance is unique but stable over time within 1% inter-annual variability
  The GEO calibration is verified with ray-matching and found to be within 0.7%
  The uncertainty maybe overestimated due to MODIS scan angle dependencies
  This allows historical and future GEO satellites with similar channels to be calibrated with multiple references
As the DCC invariant target is more accurately characterized by wavelength, diurnally, seasonally, and regionally the 1% uncertainty can be reduced