## Slide 1: GOES-16 ABI Lunar Data Preparation to GIRO
Fangfang Yu1, Xiangqian Wu2, and Xi Shao1
GOES-16 is not declared operational yet. The ABI data are currently experimental and still undergoing testing.
1: ERT, Inc@NOAA/NESDIS/STAR
2: NOAA/NESDIS/Center for Satellite Application and Research (STAR)

## Slide 2: GOES-16
First NOAA Geostationary Operational Environmental satellite in the GOES-R Series
  Launched on November 19, 2016
  Key payload: Advanced Baseline Imager (ABI)
  Locations:
    Parked at 89.5W for PLT/PLPT tests.
    Moved from 89.5W to 89.3W between August 8, 2017 and Sept. 3, 2017.
    Plan to move from 89.3W to 75.2W for operation in December 2017.

## Slide 3: Normal Lunar Scans
MESO scan in mode 3 is used to collect the Moon data.
Each MESO scan consists of two swaths
The Moon is scanned with one of the two swaths
ABI Moon data is in Level1Alpha format – calibrated but not navigated data
G16 ABI uses FPM Side 2 detectors
ABI Mode 3 timeline
Band name | Central Wavelength (µm) | EW ASD (µrad) | NS ASD (µrad) | #Det per Column | #Det per Swath
B01 | 0.47 | 22 | 22.9 | 3 | 676
B02 | 0.64 | 11 | 10.5 | 3 | 1460
B03 | 0.87 | 22 | 22.9 | 3 | 676
B04 | 1.38 | 44 | 42.0 | 6 | 372
B05 | 1.61 | 22 | 22.9 | 6 | 676
B06 | 2.25 | 44 | 42.0 | 6 | 372
ASD: Angular sample distance
swath
Best Detector Selection (BDS)

## Slide 4: Normal Lunar Scans: Chasing and Trending
Chasing:
  ABI consecutively scans the Moon as it transits across the space within the Field of Regard (FOR).
  Used for the scan mirror response vs. scan-angle (RVS) study, as well as trending of instrument degradation analysis.
  Mainly conducted during the PLT/PLPT periods.
Trending:
  ABI scans the Moon during the meso frame time which is ~2 minutes.
  Used for the monitoring and trending of the solar calibration result.
  Will be conducted 2-4 times per month throughout the satellite mission life.
Phase angle range: -90o – 90o

## Slide 5: Moon Chasing Images
Earth

## Slide 6: Moon Chasing Images
Earth

## Slide 7: Moon Chasing Images
Earth

## Slide 8: Moon Chasing Images
Earth

## Slide 9: GOES-16 ABI Lunar Images in L1Alpha Data
B1 (047)
B2 (064)
B3 (087)
B4 (138)
B5 (161)
B6 (225)
Image processing for the trending analysis
Lunar edge is >0.002 radians away from the FOR boundaries
Lunar edge is >0.013 radians away from the Earth limb
Subset boundary >0.002 radians away from the illuminated moon edge

## Slide 10: GOES-16 ABI VNIR Radiometric Calibration
Two-point on-orbit calibration targets:
  Spectralon Solar Diffuser
  Deep space
Quadratic detector responsivity correction
  Quadratic term (q) determined on-ground
  Linear calibration coefficient (m) is determined on-orbit with SD and deep space observations
Dalta et al. 2016, Comparison of the calibration algorithms and SI Traceability of MODIS, VIIRS, GOES and GOES-R ABI Sensors, Remote Sensing, 8(126), doi:10.3390/rs8020126, http://www.mdpi.com/2072-4292/8/2/126
Where fint,ch is the integration factor for viewing SCT. For ABI fint,ch=9

## Slide 11: ABI Lunar Irradiance (Currently)
ABI Irradiance:
Ω: Sample solid angle = Sr(EW_ASD * NS_ASD)=EW_ASD*NS_ASD
Radiancei,j : calibrated radiance at (i,j) image coordinate
Band name | Central Wavelength (µm) | EW ASD (µrad) | NS ASD (µrad)
B01 | 0.47 | 22 | 22.9
B02 | 0.64 | 11 | 10.5
B03 | 0.87 | 22 | 22.9
B04 | 1.38 | 44 | 42.0
B05 | 1.61 | 22 | 22.9
B06 | 2.25 | 44 | 42.0

## Slide 12: Oversampling Factor – Value by Vendor
Large under-sampling
Large over-sampling
ASD: Angular sample distance, BDS: Best detector selection, RDP: relative detector position
Harris report in EUMETSAT 2015, Toulouse, France

## Slide 13: Oversampling Factor – Proposed Calculation
Currently proposed for ABI, need further discussions within NOAA and with JMA:
ASD: Angular sample distance, BDS: Best detector selection, RDP: relative detector position
Mean angular sample distance at NS direction. Derived from RDP map for the BDS detectors
Scan rate of the EW scan mirror.
It is very stable and constant for ABI
Frame time between two neighboring samples
Determine the image spatial resolution

## Slide 14: Empirical Method
where Dmoon-sat is the distance between the Moon and GOES satellite, Ω is the sample solid angle, m is the number of illuminated moon samples, rmoon is the moon radius (1735.5km), and fractionmoon is the fraction of illuminated moon, and φis the phase angle.
Key number!
Uncertainty mainly comes from the threshold to identify the illuminated lunar images.
Blurred moon edges affect the result
Brighter moon preferred
Considering the illuminated moon area as a pseudo-pixel, the mean oversampling of this pseudo-pixel can be calculated:

## Slide 15: Oversampling Factor Values
Band# | Oversampling Factor (Theoretical) | Oversampling Factor (Empirical) | Difference (%)
B01 | 1.0126270 | 1.0134626 | 0.08
B02 | 1.0135615 | 1.0010393 | 1.24
B03 | 1.0125714 | 1.0148494 | 0.22
B04 | 1.0111070 | 1.0110186 | 0.01
B05 | 1.0127543 | 1.0258920 | 1.30
B06 | 1.0117681 | 1.0238764 | 1.12

## Slide 16: Summary
Oversampling factor may not be important for trending the instrument degradation when it is a constant.
Important for irradiance-based absolute calibration, sensor-to-sensor and band-to-band inter-calibrations
  Also important for accurate MTF analysis (Xi Shao’s report)
  For ABI-like instruments, it can be calculated with scan rate, sample frame time and RDP values
  Validated with Moon observations with near full moon images
Oversampling factor values for ABI bands are relatively small