## Slide 1
S. Wagner, B. Fougnie, T. Hewison – EUMETSAT
Inter-band calibration using the Moon

## Slide 2: Inter-band calibration
  Inter-band calibration = way to validate/monitor the relative radiometric consistency between bands of a same instrument
  A way to transfer absolute calibration from a reference band to others
   In particular for the calibration of absorption bands
  Critical in the absence of on-board calibrators
  An efficient monitoring method for on-boards calibrators
  More and more important with the use of multi-spectral information to derive Level-2 products: ocean colour, aerosols, clouds properties, etc.
   Not only Low-Earth Orbit instruments, but new generations of GEO imagers with more reflective solar bands

## Slide 3: Using the Moon for inter-band calibration
    Lunar reflectance spectrum is smooth, with only wide, shallow absorption features
    Many different spectral response functions of instruments can be accommodated
    Out-of-Atmosphere Smooth Reference spectrum:
     Convenient for inter-band calibration
     Convenient for the calibration of absorption bands

## Slide 4: Inter-band radiometric calibration
Interband normalization
Full image
Region of
interest
B. Fougnie,
Pers. Comm., 2014
Tested on PARASOL (IFOV: 6.5 km)

## Slide 5: Inter-band radiometric calibration
    Inter-comparison with other instruments
Applied successfully to MODIS and VIIRS
B. Fougnie,
GSICS annual meeting, 2016

## Slide 6: Inter-band calibration using the GLOD…
  The GLOD includes data from many LEO instruments + a few GEO instruments with more Reflective Solar Bands
  Next slides = an abrupt inter-band comparison with 2 reference bands:
    0.6 µm
    0.865 µm
  Based on data as provided to EUMETSAT. May not reflect the latest version
  Results still to be refined and discussed with data owners + LC Community
  But potential is there!
   Please, do not take those results as end results. It is a base for discussion.

## Slide 7: An example using the data GLOD – 0.6 µm
VIIRS NASA
Results for FY3C are not correct due to the wrong selection of the reference channel

## Slide 8: An example using the data GLOD – 0.6 µm
VIIRS NASA
Results for FY3C are not correct due to the wrong selection of the reference channel

## Slide 9: An example using the data GLOD – 0.8 µm
VIIRS NASA
Results for FY3C are not correct due to the wrong selection of the reference channel

## Slide 10: An example using the data GLOD – 0.8 µm
VIIRS NASA

## Slide 11: An example using the data GLOD – 0.865 µm
VIIRS NASA
SCIAMACHY at 50° phase
Results for FY3C are not correct due to the wrong selection of the reference channel

## Slide 12: An example using the data GLOD – 0.865 µm
VIIRS NASA
SCIAMACHY at 50° phase

## Slide 13: Meteosat-10 - SEVIRI inter-band – 0.6 µm
VIIRS NASA
GIRO
Variability due to the changing illumination conditions

## Slide 14: Meteosat-10 - SEVIRI inter-band – 0.865 µm
VIIRS NASA
GIRO

## Slide 15: Himawari-8 - AHI inter-band – 0.6 µm
VIIRS NASA
GIRO

## Slide 16: Himawari-8 - AHI inter-band – 0.865 µm
VIIRS NASA
GIRO

## Slide 17: Conclusion
More work required on the current results (SBAFs, reference bands selection, etc.)
However, it is an efficient method to investigate the consistency between:
  Inter-band calibration methods
  Absolute calibration
  Consistency between instruments
A way to check the Moon irradiance calculation when comparing with other methods???
Would need the teams interested to provide more data (up-to-date data processing)
Interactions needed with the instrument experts and data owners to define the reference band(s)