## Slide 1
Tim Hewison
GSICS IR Reference Uncertainty & Traceability Report

## Slide 2: IR Reference Sensor Traceability & Uncertainty Report
Aims
  To support choice of reference instruments for GSICS and Metop-A/IASI as Anchor
  To provide traceability between reference instruments (IASI, AIRS, CrIS)
  By consolidating pre-launch test results and various in-flight comparisons
  To seek consensus on uncertainties in absolute calibration of reference sensors
Limitations
  No new results, just expressing results of existing comparisons in a common way,
  reformatting where necessary, to allow easy comparisons
Error Budget & Traceability
  Focus on radiometric and spectral calibration – for AIRS, IASI, CrIS
Inter-comparisons
  Introduction: Pros and Cons of each method
  Direct Comparisons: Polar SNOs, Tandem SNOs (AIRS+CrIS), Quasi-SNOs,
  Double-Differencing: GEO-LEO, NWP+RTM, Aircraft campaigns
  Other Methods: Regional Averages (“Massive Means”), Reference Sites (Dome-C..)

## Slide 3: IR Reference Sensor Inter-Comparisons
 |  |  |  |  | Mean Difference dTb [K] |  |  |  |  | 
 |  | Pseudo Channel [cm-1] | Min Freq [cm-1] | Max Freq [cm-1] | 200 | 220 | 240 | 260 | 280 | 300
Start date |  | 655 | 650 | 660 |  |  |  |  |  | 
End date |  | 665 | 660 | 670 |  |  |  |  |  | 
Start time |  | 675 | 670 | 680 |  |  |  |  |  | 
End time |  | 685 | 680 | 690 |  |  |  |  |  | 
Min Latitude [°] |  | 695 | 690 | 700 |  |  |  |  |  | 
Max Latitude [°] |  | 705 | 700 | 710 |  |  |  |  |  | 
Min Longitude [°] |  | 715 | 710 | 720 |  |  |  |  |  | 
Max Longitude [°] |  | 725 | 720 | 730 |  |  |  |  |  | 
Min Scan Angle [°] |  | 735 | 730 | 740 |  |  |  |  |  | 
Max Scan Angle [°] |  | 745 | 740 | 750 |  |  |  |  |  | 
 |  | 755 | 750 | 760 |  |  |  |  |  | 
Collocation method: |  | 765 | 760 | 770 |  |  |  |  |  | 
Collocation dist [km] |  | 775 | 770 | 780 |  |  |  |  |  | 
Collocation time [s] |  | 785 | 780 | 790 |  |  |  |  |  | 
Collocation sec(theta) |  | 795 | 790 | 800 |  |  |  |  |  | 
Filtering applied |  | 805 | 800 | 810 |  |  |  |  |  | 
 |  | 815 | 810 | 820 |  |  |  |  |  | 
Algorithm Ref |  | 825 | 820 | 830 |  |  |  |  |  | 
Dataset Ref |  | 835 | 830 | 840 |  |  |  |  |  | 
Monitored Instrument |  | 845 | 840 | 850 |  |  |  |  |  | 
Processing Version |  | 855 | 850 | 860 |  |  |  |  |  | 
Dataset Ref |  | 865 | 860 | 870 |  |  |  |  |  | 
Reference Instrumemt |  | 875 | 870 | 880 |  |  |  |  |  | 
Processing Version |  | 885 | 880 | 890 |  |  |  |  |  | 
Dataset Ref |  | 895 | 890 | 900 |  |  |  |  |  | 
Form consensus on relative calibration
  Re-binning results of existing comparisons
  to make them comparable:
Biases with respect to Metop-A/IASI
  With standard uncertainties (k=1)
At full spectral resolution
  In CrIS channel-space
  – or in 10cm-1 bins within AIRS bands
  Averaged over specific spectral bands
  Or average results over broad-band channels
  With specific SRFs - rectangular?
Converted into Brightness Temperatures
  For specific radiance scenes
  i.e. 200K, 210K, … 300K
For all viewing angles
  and/or for specific ranges - e.g. nadir ±10°
Over specific period - e.g. at least 1 year
  Common 3 year period from IASI-B start

## Slide 4: Summary of Previous Web Meeting (2016-06-21)
The proposed structure of the report was agreed,
  with the addition of a sub-section in the introduction to address the need for continuous monitoring of the reference instruments' calibration.
Additional sub-sub-sections were also identified to briefly address
  a) radiometric noise,
  b) spectral calibration and
  c) geometric factors (navigation accuracy etc) in the error budget
Although these need not be treated in a fully rigorous approach,
  given their negligible impact on the inter-calibration products [for a) and b)] and
  the difficulty of assessment [for c)].
The contributor authors to each sub-section were identified
  - either as firm, or tentative.
The spectral resolution of the comparisons was discussed at length and different spectral conversion methods described.
It was felt that 10cm-1 bins would be sufficient.
It seems the most difficult issue is dealing with AIRS' gap channels.
It was agreed that further discussion on this topic is needed,
  so another web meeting will be set up to discuss this in mid-August 2016.

## Slide 5: Action GRWG.20160621.1
Action GRWG.20160621.1: Tim Hewison (EUMETSAT) to check with NIST/NPL and confirm the recommended coverage factor to be used for error budgets and comparisons.
  Action completed 2016-08-03, with the following response from Emma Woolliams (NPL) - and agreed by Dave Walker (NIST):
  “The uncertainty analysis should all be performed with standard uncertainties.
  Any uncertainty budget (table) should definitely be full of standard uncertainties.
  The adding in quadrature (applying the Law of Propagation of Uncertainties) must be done with standard uncertainties.
  But the final result may be quoted as an expanded uncertainty.
  In which case the k value must be provided and if it’s not 2, the number of degrees of freedom should be provided too.
  That means that other people can divide by the right number when including your uncertainty analysis into their budgets.”

## Slide 6: Action GRWG.20160621.2
Action GRWG.20160621.2: Denis Jouglet (CNES) to distribute spectral averaging coefficients and documentation describing their application by early July. - Action completed 2016-06-21 - See next slide sent by email.

## Slide 7: IASI / AIRS : methodology – Denis Jouglet
Spectral match:
  Work with IASI L1C, AIRS L1B
  Method: 33 broad pseudo-bands (PBs) from GSICS
  1 PB = summation of ~100s of elementary channels (most widths between 23 and 63 cm-1)
  Reduces noise and spectral resolution differences
  AIRS spurious channels: taken into account through a weighted summation of the IASI channels
  (weighs are computed to make the resulting PB response functions similar in IASI and AIRS)
Comparison of ΔT = TIASI - TAIRS in each PB
  Other methods under progress
    similar channels (statistical similar behavior)
    convolved channels
Instrumental functions of one PB for AIRS (including spurious channels), for IASI without weighting in the channels summation and for IASI with weighting

## Slide 8: Summary of Previous Web Meeting (2016-09-08)
Spectral averaging methods were reviewed by Denis Jouglet
  Action GIR.20160908.1: Denis Jouglet (CNES) to apply spectral averaging method to calculate static weightings for generating 10cm-1 pseudo channels for AIRS-IASI comparison over 3 year period (2013-03-01/2016-03-01) - and consider application for CrIS -IASI comparisons.
Inter-comparison database was introduced by Tim Hewison
  It was agreed that the proposed 10cm^-1 spectral binning is adequate
  It was agreed that finer radiance binning is needed to ensure results are comparable (linear)
  Action GRWG.20160908.3: Dave Tobin (SSEC) to regenerate comparison results in 10K bins over 3 year period (2013-03-01/2016-03-01), describe method and share raw SNO results.
  Action GRWG.20160908.4: Tim Hewison (EUMETSAT) to regenerate comparison results in 10K bins and redo double-difference analysis, expressing results in BT, radiance and % radiance – done see next slides.
Comparisons of AATSR and IASI were introduced by Manik Bali
  Recommendation: Manik Bali (NOAA) to investigate adding incidence angle matching to AATSR-IASI comparison, with weighting according to the variance of the SNO radiances.
  Recommendation: Manik Bali (NOAA) to review outline for report on Traceability and Uncertainty of GSICS Infrared Reference Sensors and propose how his AATSR-IASI analysis could be included/referenced.

## Slide 9: Summary of Previous Web Meeting (2017-06-20) # 1
Tim Hewison provided an update on the (slow) progress with the Report on "Traceability and Uncertainty of GSICS Infrared Reference Sensors".
It was agreed to report the uncertainties and relative differences between sensors as brightness temperatures differences,
evaluated for scenes ranging from 200-300K at 10K intervals,
and at 10cm^-1 intervals, calculated from hyperspectral radiances, following the methodology previously described by Denis Jouglet.
It was further suggested that the acronym "IRRefUTable" be informally adopted as a shorthand to refer to activities related to this report.

## Slide 10: Summary of Previous Web Meeting (2017-06-20) #2
Elsa Jacquette described the process of revising the IASI uncertainty budget, which is planned for this year. She will circulate a draft by September, for inclusion as part of the planned presentation on the "IRRefUTable Report": https://iasi.cnes.fr/sites/default/files/drupal/201612/default/bpc_iasi-conference4-1_03_non_linearity_jacquette.pdf
Joe Taylor confirmed that Dave Tobin's recent updates to the counterpart analysis for CrIS still follow the same basic approach as described in his paper, which should serve as an example for AIRS and IASI - although some of the values have been updated.
Traceability to SI standards was discussed.
  Although it is possible to establish this during pre-launch testing, the traceability chain is broken during launch and is not possible to re-establish with the current instruments.
  Furthermore, it is difficult to quantify the uncertainties associated with all the changes since pre-launch testing.
  However, the IRRefUTable report attempts to establish IASI, AIRS and CrIS as Community Reference Standards.

## Slide 11: Summary of Previous Web Meeting (2017-06-20) #3
Tom Pagano presented a draft of the v7 of the AIRS calibration. This is the first update since launch.
The biggest change is in the treatment of the scan mirror reflectivity-polarisation interaction.
The observations in the 4 space views have been analysed over the lifetime of AIRS to derive a time-dependent correction.
The overall uncertainty of the calibration is reduced by a factor of ~2 (at 220K) in most bands.
Tom will prepare some slides for inclusion as part of the planned presentation on the "IRRefUTable Report"

## Slide 12: Summary of Previous Web Meeting (2017-06-20) #4
Double differencing of IASI-A/B against geostationary imagers' IR Channels
  Tim Hewison (EUMETSAT) - Meteosat imagers
  Masaya Takahashi (JMA) - Himawari-8 imager (AHI)
  Fangfang Yu (NOAA) – GOES imagers
  Hanlie Xu or Na Xu (CMA) - FY2/4 imager
  Minju Gu (KMA) - COMS imager
Igor Tomazic (EUMETSAT) – Double differencing of IASI-A/B against SLSTR

## Slide 13: 2013-03/2017-03 (SEVIRI-IASIA)-(SEVIRI-IASIB) - Tb
 |  |  | Mean Difference dTb [K] |  |  |  |  |  |  |  |  |  |  | 
Channel | 50% [cm-1] | 50% [cm-1] | 200 | 210 | 220 | 230 | 240 | 250 | 260 | 270 | 280 | 290 | 300 | K
IR13.4 | 714 | 782 | -0.30 | -0.24 | -0.19 | -0.15 | -0.12 | -0.09 | -0.06 | -0.04 | -0.02 | 0.00 | 0.02 | K
IR12.0 | 800 | 870 | -0.16 | -0.13 | -0.11 | -0.09 | -0.07 | -0.06 | -0.05 | -0.04 | -0.04 | -0.03 | -0.02 | K
IR10.8 | 885 | 971 | -0.27 | -0.21 | -0.16 | -0.13 | -0.10 | -0.08 | -0.06 | -0.05 | -0.04 | -0.03 | -0.02 | K
IR9.7 | 1018 | 1047 | -0.19 | -0.14 | -0.11 | -0.08 | -0.06 | -0.05 | -0.04 | -0.03 | -0.02 | -0.01 | -0.01 | K
IR8.7 | 1124 | 1177 | -0.09 | -0.07 | -0.05 | -0.04 | -0.03 | -0.02 | -0.01 | -0.01 | -0.01 | 0.00 | 0.00 | K
IR7.3 | 1316 | 1409 | 0.01 | 0.00 | 0.00 | 0.00 | 0.00 | 0.00 | 0.00 | 0.00 | 0.00 | 0.00 | 0.00 | K
IR6.2 | 1493 | 1724 | -0.14 | -0.08 | -0.05 | -0.03 | -0.01 | 0.00 | 0.01 | 0.01 | 0.02 | 0.02 | 0.03 | K
IR3.9 | 2385 | 2751 | 0.02 | 0.01 | 0.00 | 0.00 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | K
Chart: 
| 200 | 210 | 220 | 230 | 240 | 250 | 260 | 270 | 280 | 290 | 300
IR13.3 | -0.3 | -0.24 | -0.19 | -0.15 | -0.12 | -0.09 | -0.06 | -0.04 | -0.02 | 0 | 0.02
IR11.9 | -0.16 | -0.13 | -0.11 | -0.09 | -0.07 | -0.06 | -0.05 | -0.04 | -0.04 | -0.03 | -0.02
IR10.8 | -0.27 | -0.21 | -0.16 | -0.13 | -0.1 | -0.08 | -0.06 | -0.05 | -0.04 | -0.03 | -0.02
IR09.7 | -0.19 | -0.14 | -0.11 | -0.08 | -0.06 | -0.05 | -0.04 | -0.03 | -0.02 | -0.01 | -0.01
IR08.7 | -0.09 | -0.07 | -0.05 | -0.04 | -0.03 | -0.02 | -0.01 | -0.01 | -0.01 | 0 | 0
IR07.4 | 0.01 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
IR06.3 | -0.14 | -0.08 | -0.05 | -0.03 | -0.01 | 0 | 0.01 | 0.01 | 0.02 | 0.02 | 0.03
IR03.9 | 0.02 | 0.01 | 0 | 0 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01 | -0.01
  NWP Double-Differencers:

## Slide 14
Mean Difference dTb [K]
Uncertainty on Mean Difference u(dTb) [K], k=1
IASI-B – IASI-A Stats – Inter-comparison between AHI and SEVIRI
Masaya Takahashi (JMA)
Color: Spectral Bands of AHI and SEVIRI
Himawari-8/AHI
(2015-07-01 – 2017-06-30)
Meteosat-10/SEVIRI
(2013-03-01 – 2017-03-01)
Meteosat-10/SEVIRI
(2013-03-01 – 2017-03-01)
Himawari-8/AHI
(2015-07-01 – 2017-06-30)
Good agreements w/ SEVIRI double difference results except 3.9 m

## Slide 15: Way Forward – Report
Report and “Database” initially in Google Docs – now switched to Zoho
Confirm Contributors:
  Error Budget & Traceability
    AIRS – Tom Pagano
    CrIS – Dave Tobin
    IASI – Denis Jouglet
  Polar SNOs
    Denis Jouglet, Dave Tobin, Likun Wang
  Tandem SNOs
    Dave Tobin
  Quasi SNOs
    Denis Jouglet
  GEO-LEO Double Differencers:
    Tim Hewison (EUMETSAT), Masaya Takahashi (JMA), Fangfang Yu/Fred Wu (NOAA)
    Hanlie Xu / Na Xu (CMA) ?, Minju Gu (KMA) ?
  NWP Double-Differencers:
    Fred Wu (NOAA), Chengli Qi (CMA)
  Aircraft Double-Differencers:
    Joe Taylor?

## Slide 16: Way Forward – Report Timeline
Include N20/CrIS?
  Launch Nov 2017
Include Metop-C/IASI
  Launch Nov 2018 – Need ~6+ months
  IASI Error Budget ~mid 2019
Include FY-3D/HIRAS?
  Launch Nov 2017. Data from Oct 2018?
Revise Common Period?
  Previously agreed:
  3 year period (2013-03-01/2016-03-01)
Write-up Report
  2019?

## Slide 17
Thank You!