## Slide 1: Inter-comparison of NOAA-20 and Suomi-NPP ATMS Data
Ninghai SUNa, Hu “Tiger” YANGb, *Robbie IACOVAZZIa, Quanhua “Mark” LIUc
aGST, Inc. Greenbelt, MD 20770 USA
bUMD/ESSIC, College Park, MD 20740 USA
cNOAA, NESDIS, STAR, College Park, MD 20740 USA
GSICS Microwave Subgroup “Virtual” Meeting
18 March 2020
*Presenter
Notes: Acknowledge JPSS Program

## Slide 2: Outline
NOAA-20 and S-NPP Advanced Technology Microwave Sounder (ATMS) data monitoring and analysis tools
NOAA-20 and S-NPP Data Comparisons
  Instrument Noise
  Geolocation Error
  ATMS Antenna Temperature (Ta) TDR and/or Brightness Temperature (Tb) SDR Biases
Summary

## Slide 3: NOAA-20 and S-NPP ATMS data monitoring and analysis tools
Tracking and trending ATMS measurements is critical to ensuring data quality for use in NOAA numerical weather prediction and climate analysis.
NOAA has established a web-based Integrated Calibration/Validation System in part to fulfill this need (https://www.star.nesdis.noaa.gov/icvs/).
Current short- and long-term tracking an trending methods for ATMS
*Engineering, calibration, and data-quality flag parameters
*Geolocation error determined with the coastline method
Ta (TDR) and/or Tb (SDR) biases determined using various methods
*Generates NOAA-20 and S-NPP ATMS data quality comparison results.

## Slide 4: Methods to Monitor TDR and/or SDR Biases
  Observed- and Community Radiative Transfer Model (CRTM)-simulated background (O-B) Ta and/or Tb biases computed separately using NWP output and Radio-occultation soundings
  Inter-satellite simultaneous nadir overpass (SNO) Ta inter-comparisons with AMSU-A
  *Inter-satellite Ta and/or Tb 32-day running-mean difference
  *Inter-satellite O-B Ta/Tb bias double difference
*Generates NOAA-20 and S-NPP ATMS data quality comparison results.

## Slide 5: NOAA-20 and S-NPP ATMS data comparisons
Engineering, calibration, and data-quality flag parameters – Instrument Noise (Channel 1-6)
Smaller NEDT

## Slide 6: NOAA-20 and S-NPP ATMS data comparisons
Engineering, calibration, and data-quality flag parameters – Instrument Noise (Channel 7-12)
Smaller NEDT

## Slide 7: NOAA-20 and S-NPP ATMS data comparisons
Engineering, calibration, and data-quality flag parameters – Instrument Noise (Channel 13-18)
Smaller NEDT

## Slide 8: NOAA-20 and S-NPP ATMS data comparisons
Engineering, calibration, and data-quality flag parameters – Instrument Noise (Channel 19-22)
Smaller NEDT
NOAA-20 ATMS has <20% smaller noise than S-NPP
S-NPP and NOAA-20 ATMS have stable noise characteristics

## Slide 9: NOAA-20 and S-NPP ATMS data comparisons
Geolocation Error (K and Ka Bands)
Smaller Geo Error Variability but Larger Error
Smaller Geo Error Variability
Large Geo Error
Large Geo Error Variability
K Band
Ka Band

## Slide 10: NOAA-20 and S-NPP ATMS data comparisons
Geolocation Error (V and W Bands)
Large Geo Error and Error Variability
NOAA-20 ATMS often has less absolute geolocation error and error variability than S-NPP ATMS
V Band
W Band
Note - G Band is too opaque to apply coastline method

## Slide 11: NOAA-20 and S-NPP ATMS data comparisons
ATMS Ta (TDR) and/or Tb (SDR) Biases (Channel 1)
Negative bias (~-0.25 K) for both TDR and SDR
Possible seasonal cycle in 32-day running-mean global mean bias difference
SDR has a larger magnitude bias until October 15, 2019 software patch to implement reflector emission correction and NPP antenna pattern coefficients update.
TDR-SDR Bias mismatch.
Possible seasonal cycle.

## Slide 12: NOAA-20 and S-NPP ATMS data comparisons
ATMS Ta (TDR) and/or Tb (SDR) Biases (Channel 2)
Negative bias (~-0.20 K) for both TDR and SDR
Possible seasonal cycle in 32-day running-mean global mean bias difference
SDR has a larger magnitude bias until October 15, 2019 software patch to implement reflector emission correction and NPP antenna pattern coefficients update.
Possible seasonal cycle.

## Slide 13: NOAA-20 and S-NPP ATMS data comparisons
ATMS Ta (TDR) and/or Tb (SDR) Biases (Channel 3)
Negative bias (~-0.30 K) for both TDR and SDR for 32-day running mean, but double-difference of ECMWF O-B measurements shows a near zero bias.
SDR has a larger magnitude bias until October 15, 2019 software patch to implement reflector emission correction and NPP antenna pattern coefficients update.
Bias mismatch.

## Slide 14: NOAA-20 and S-NPP ATMS data comparisons
ATMS Ta (TDR) and/or Tb (SDR) Biases (Channel 4)
Negative bias (~0.50 K) for both TDR and SDR for 32-day running mean, but double-difference of ECMWF O-B measurements shows a near zero bias.
SDR has a larger magnitude bias until October 15, 2019 software patch to implement reflector emission correction and NPP antenna pattern coefficients update.
Bias mismatch.

## Slide 15: NOAA-20 and S-NPP ATMS data comparisons
ATMS Ta (TDR) and/or Tb (SDR) Biases (Channel 5)
Negative bias (~0.35 K) for 32-day running mean, but double-difference of ECMWF and RO O-B measurements show a near zero bias. Note that RO results are TDR only.
Double difference from RO has larger variance due to lower sampling and surface contamination.
SDR has a larger magnitude bias until October 15, 2019 software patch to implement reflector emission correction and NPP antenna pattern coefficients update.
Bias mismatch.
Large variance due to relatively small population and surface contamination.

## Slide 16: NOAA-20 and S-NPP ATMS data comparisons
ATMS Ta (TDR) and/or Tb (SDR) Biases (Channel 6)
Negative bias (~-0.3 K) for 32-day running mean, and double-difference of ECMWF and RO O-B measurements. Note that RO results are TDR only.
Double difference from RO has larger variance due to lower sampling and surface contamination.
SDR has a larger magnitude bias until October 15, 2019 software patch to implement reflector emission correction and NPP antenna pattern coefficients update.

## Slide 17: NOAA-20 and S-NPP ATMS data comparisons
ATMS Ta (TDR) and/or Tb (SDR) Biases (Channel 7)
Negative bias (~-0.35 K) for 32-day running mean, and double-difference of ECMWF and RO O-B measurements. Note that RO results are TDR only.
Possible seasonal cycle in 32-day running-mean global mean bias difference
Double difference from RO has larger variance due to lower sampling.
SDR has a larger magnitude bias until October 15, 2019 software patch to implement reflector emission correction and NPP antenna pattern coefficients update.
Possible seasonal cycle.

## Slide 18: NOAA-20 and S-NPP ATMS data comparisons
ATMS Ta (TDR) and/or Tb (SDR) Biases (Channel 8)
Negative bias (~-0.2 K) for 32-day running mean, and double-difference of ECMWF and RO O-B measurements. Note that RO results are TDR only.
Possible seasonal cycle in 32-day running-mean global mean bias difference
Double difference from RO has larger variance due to lower sampling.
SDR has a larger magnitude bias until October 15, 2019 software patch to implement reflector emission correction and NPP antenna pattern coefficients update.

## Slide 19: NOAA-20 and S-NPP ATMS data comparisons
ATMS Ta (TDR) and/or Tb (SDR) Biases (Channel 9)
Negative bias (~-0.35 K) for 32-day running mean, and double-difference of ECMWF and RO O-B measurements. Note that RO results are TDR only.
Possible seasonal cycle in 32-day running-mean global mean bias difference
Double difference from RO has larger variance due to lower sampling.
SDR has a larger magnitude bias until October 15, 2019 software patch to implement reflector emission correction and NPP antenna pattern coefficients update.

## Slide 20: NOAA-20 and S-NPP ATMS data comparisons
ATMS Ta (TDR) and/or Tb (SDR) Biases (Channel 10)
Negative bias (~-0.25 K) for 32-day running mean, and double-difference of ECMWF and RO O-B measurements. Note that RO results are TDR only.
Possible seasonal cycle in 32-day running-mean global mean bias difference
Double difference from RO has larger variance due to lower sampling.
SDR has a larger magnitude bias until October 15, 2019 software patch to implement reflector emission correction and NPP antenna pattern coefficients update.

## Slide 21: NOAA-20 and S-NPP ATMS data comparisons
ATMS Ta (TDR) and/or Tb (SDR) Biases (Channel 11)
Negative bias (~-0.15 K) for 32-day running mean, and double-difference of ECMWF and RO O-B measurements. Note that RO results are TDR only.
Possible seasonal cycle in 32-day running-mean global mean bias difference
Double difference from RO has larger variance due to lower sampling.
SDR has a larger magnitude bias until October 15, 2019 software patch to implement reflector emission correction and NPP antenna pattern coefficients update.

## Slide 22: NOAA-20 and S-NPP ATMS data comparisons
ATMS Ta (TDR) and/or Tb (SDR) Biases (Channel 12)
Negative bias (~-0.15 K) for 32-day running mean, and double-difference of ECMWF and RO O-B measurements. Note that RO results are TDR only.
Possible seasonal cycle in 32-day running-mean global mean bias difference
Double difference from RO has larger variance due to lower sampling.
SDR has a larger magnitude bias until October 15, 2019 software patch to implement reflector emission correction and NPP antenna pattern coefficients update.

## Slide 23: NOAA-20 and S-NPP ATMS data comparisons
ATMS Ta (TDR) and/or Tb (SDR) Biases (Channel 13)
Negative bias (~-0.30 K) for 32-day running mean, and double-difference of ECMWF and RO O-B measurements. Note that RO results are TDR only.
Possible seasonal cycle in 32-day running-mean global mean bias difference
Double difference from RO has larger variance due to lower sampling.
SDR has a larger magnitude bias until October 15, 2019 software patch to implement reflector emission correction and NPP antenna pattern coefficients update.

## Slide 24: NOAA-20 and S-NPP ATMS data comparisons
ATMS Ta (TDR) and/or Tb (SDR) Biases (Channel 14)
Negative bias (~-0.30 K) for both TDR and SDR
Possible seasonal cycle in 32-day running-mean global mean bias difference
SDR has a larger magnitude bias until October 15, 2019 software patch to implement reflector emission correction and NPP antenna pattern coefficients update.

## Slide 25: NOAA-20 and S-NPP ATMS data comparisons
ATMS Ta (TDR) and/or Tb (SDR) Biases (Channel 15)
Near zero bias for both TDR and SDR
Possible seasonal cycle in 32-day running-mean global mean bias difference
SDR has a larger magnitude bias until October 15, 2019 software patch to implement reflector emission correction and NPP antenna pattern coefficients update.

## Slide 26: NOAA-20 and S-NPP ATMS data comparisons
ATMS Ta (TDR) and/or Tb (SDR) Biases (Channel 16)
Negative bias (~-0.15 K) for 32-day running mean, double-difference of ECMWF O-B measurements shows a 0.15 K positive bias.
Possible seasonal cycle in 32-day running-mean global mean bias difference
SDR has a larger magnitude bias until October 15, 2019 software patch to implement reflector emission correction and NPP antenna pattern coefficients update.

## Slide 27: NOAA-20 and S-NPP ATMS data comparisons
ATMS Ta (TDR) and/or Tb (SDR) Biases (Channel 17)
Negative bias (~-0.90 K) for 32-day running mean, double-difference of ECMWF O-B measurements shows a near zero bias.
Possible seasonal cycle in 32-day running-mean global mean bias difference
SDR has a larger magnitude bias until October 15, 2019 software patch to implement reflector emission correction and NPP antenna pattern coefficients update.
Substantial bias mismatch.

## Slide 28: NOAA-20 and S-NPP ATMS data comparisons
ATMS Ta (TDR) and/or Tb (SDR) Biases (Channel 18)
Negative bias (~-0.80 K) for 32-day running mean, double-difference of ECMWF O-B measurements shows a bias of about -0.3.
SDR has a larger magnitude bias until October 15, 2019 software patch to implement reflector emission correction and NPP antenna pattern coefficients update.

## Slide 29: NOAA-20 and S-NPP ATMS data comparisons
ATMS Ta (TDR) and/or Tb (SDR) Biases (Channel 19)
Negative bias (~-0.70 K) for 32-day running mean, double-difference of ECMWF O-B measurements shows a bias of about -0.3.
SDR has a larger magnitude bias until October 15, 2019 software patch to implement reflector emission correction and NPP antenna pattern coefficients update.

## Slide 30: NOAA-20 and S-NPP ATMS data comparisons
ATMS Ta (TDR) and/or Tb (SDR) Biases (Channel 20)
Negative bias (~-0.50 K) for 32-day running mean, double-difference of ECMWF O-B measurements shows a bias of about -0.3.
SDR has a larger magnitude bias until October 15, 2019 software patch to implement reflector emission correction and NPP antenna pattern coefficients update.

## Slide 31: NOAA-20 and S-NPP ATMS data comparisons
ATMS Ta (TDR) and/or Tb (SDR) Biases (Channel 21)
Negative bias (~-0.40 K) for 32-day running mean, double-difference of ECMWF O-B measurements shows a bias of about -0.2.
SDR has a larger magnitude bias until October 15, 2019 software patch to implement reflector emission correction and NPP antenna pattern coefficients update.

## Slide 32: NOAA-20 and S-NPP ATMS data comparisons
ATMS Ta (TDR) and/or Tb (SDR) Biases (Channel 22)
Negative bias (~-0.30 K) for 32-day running mean, double-difference of ECMWF O-B measurements shows a bias of about -0.2.
SDR has a larger magnitude bias until October 15, 2019 software patch to implement reflector emission correction and NPP antenna pattern coefficients update.

## Slide 33
Summary
Tracking and trending ATMS measurements is critical to ensuring data quality for use in NOAA numerical weather prediction and climate analysis.
NOAA-20 and S-NPP ATMS performance are compared using
Engineering, calibration, and data-quality flag parameters
Geolocation error determined with the coastline method
Inter-satellite Ta and/or Tb 32-day running-mean difference, or O-B Ta/Tb bias double difference
NOAA-20 and S-NPP ATMS Noise Comparison
NOAA-20 ATMS has <20% smaller noise than S-NPP
S-NPP and NOAA-20 ATMS have stable noise characteristics
NOAA-20 and S-NPP ATMS Geolocation Error Comparison
NOAA-20 ATMS often has less absolute geolocation error and error variability than S-NPP ATMS
NOAA-20 and S-NPP ATMS TDR and/or SDR Comparison
SDR has a larger magnitude bias until October 15, 2019 software patch to implement reflector emission correction and S-NPP antenna pattern coefficients update.
Biases are usually negative, and vary between about -0.15K and -0.9 K.
There can be large bias differences between the 32-day running mean difference, and those determined from the double-difference of ECMWF O-B measurements. Note that RO results are TDR only.
Possible seasonal cycle in 32-day running-mean global mean bias difference for many channels
Double difference from RO has larger variance due to lower sampling, and for some channels surface contamination.
Acknowledgements: Supported by the JPSS Program
Notes: 33
33