## Slide 1
Donglin Meng
mengdl@nim.ac.cn
National Institute of Metrology (NIM), P. R. China
16 February 2021
Metrology of Antenna
Radiation Pattern
GSICS Microwave Subgroup Meeting

## Slide 2: Outlines
1
Background & Problem
NIM Spherical near field facility
3
Uncertainty evaluations
4
Methods for Calibrating antenna radiation pattern
2

## Slide 3: Comparability of remote-sensing measurements ~reproducibility , repeatability
Provided by Jieying He from National Space Science Center, CAS
Comparability: different remote-sensing measurements by different instruments at different times can be compared;
Traceability： accurate & adequate calibrations.

## Slide 4: Antenna radiation pattern ~ Visibility function
[1] Corbella, N. Duffo, M. Vall-llossera, A. Camps, and F. Torres, “The visibility function in interferometric aperture synthesis radiometry,”IEEE Trans. Geosci. Remote Sens., vol. 42, no. 8, pp. 1677–1682, Aug. 2004.
Visibility function equation -Interferometric Aperture Synthesis Radiometry
Normalized antenna radiation pattern
Maximium directivity
How to calibrate radiation pattern
How to evaluatte the uncertainty?
Notes: Polarimetric formulation of the visibility function equation including cross-polar antenna patterns
Published in: IEEE Geoscience and Remote Sensing Letters ( Volume: 2, Issue: 3, July 2005)
Page(s): 292 - 295
Date of Publication: 18 July 2005
ISSN Information:
INSPEC Accession Number: 8592729
DOI: 10.1109/LGRS.2005.846885

## Slide 5
Antenna radiation pattern——Comprehensive parameter
Antenna
A transceiver between guided wave and free-space E.M. fields.
E.M. fields in free-space
Guideded Wave
Source /Receiver
Antenna radiation pattern:
The spatial distribution of a quantity that characterizes the electromagnetic field generated by an antenna.
3D pattern
Phase center
Half-power bandwidth
Sidelobe level

## Slide 6: Antenna calibrating capability in NIM
TEM/GTEM/µTEM/ECSM
（环天线, 鞭天线）
开阔试验场 OATS (60m×40m)
远场（EMC天线）
平面近场 PNF
外推法（Extrapolation range）
球面近场 SNF

## Slide 7: Comparison of antenna calibration capability
Organization | Country | Facility | Frequency（GHz） | Uncertainty of Sidelobe level dB | Notes
NIST | USA | Roll / Azimuth Robot arms | 2 ~ 60 50 ~ 110 100 ~ 500 | Down up to -50 dB or - 60 dB； Axial ratio measurement uncertainty: 0.1 dB/dB | （23±1）℃
NPL | UK | Integrated with three antenna extrapolation rage, Roll / Azimith | 0.4 – 110 ? | Similar to NIM | （23±1）℃
KRISS | Korean | Planar/Cylind/Spherical | 2 ～ 110 ? | U=3.44 dB @ - 40dB (k=1) | （23±1）℃
TUD | Denmark | Roll / Azimuth Robot arms | 400MHz ~ 50 GHz* | @5.405GHz U (D)= 0.03dB @ HPBW | The most famous one in the world
NIM | China | Integrated with three antenna extrapolation rage, Roll / Azimith | 0.4 - 110 | 400MHz~490MHz : -40dB: U =12.9dB (k=2) 2.6 GHz~110GHz -40dB U =(1.3 ~1.8) dB (k=2) | （23±1）℃

## Slide 8: Outlines
2
Methods for Calibrating antenna radiation pattern
NIM Spherical near-field facility
3
Uncertainty evaluation
4
Background & Problem
1

## Slide 9: Basic Calibration methods
Compact range
Near-field
Far-field method
Methods
Cylindal
near field
Spherical near field
Planer near field
G >15 dBi
Near far field
Plane wave generator
Large antennas, economical

## Slide 10: Challenges for calibrating accurately radiation pattern
Far-field:
R= 4.5 D2/λ
Far field distance
Low sidelobe level
Phase center
Calculated radiation patterns of a paraboloid antenna for different distances from the antenna， 1970.
Far field
Spherical near field
Notes: Calculated radiation patterns of a paraboloid antenna for different distances from the antenna. (SOURCE: J. S. Hollis, T. J. Lyon, and L. Clayton, Jr. (eds.), Microwave Antenna Measurements, Scientific-Atlanta, Inc., July 1970).

## Slide 11: Far field distance
Chart: Far field distance
| 1 | 2 | 5 | 10 | 20 | 40 | 80 | 160 | 320 | 640 | 1280
Aperture 0.2 m | 1.3333 | 2.6667 | 6.6667 | 13.3333 | 26.6667 | 53.3333 | 106.6667 | 213.3333 | 426.6667 | 853.3333 | 1706.6667
Aperture 0.5 m | 3.3333 | 6.6667 | 16.6667 | 33.3333 | 66.6667 | 133.3333 | 266.6667 | 533.3333 | 1066.6667 | 2133.3333 | 4266.6667
Aperture 1 m | 6.6667 | 13.3333 | 33.3333 | 66.6667 | 133.3333 | 266.6667 | 533.3333 | 1066.6667 | 2133.3333 | 4266.6667 | 8533.3333
Aperture 2 m | 13.3333 | 26.6667 | 66.6667 | 133.3333 | 266.6667 | 533.3333 | 1066.6667 | 2133.3333 | 4266.6667 | 8533.3333 | 17066.6667
Aperture 5 m | 33.3333 | 66.6667 | 166.6667 | 333.3333 | 666.6667 | 1333.3333 | 2666.6667 | 5333.3333 | 10666.6667 | 21333.3333 | 42666.6667
Aperture 7 m | 46.6667 | 93.3333 | 233.3333 | 466.6667 | 933.3333 | 1866.6667 | 3733.3333 | 7466.6667 | 14933.3333 | 29866.6667 | 59733.3333

## Slide 12: Planar near-field (PNF) requirements
Measurement distance:
Measurement area:
D is the aperture of the antenna under test.

## Slide 13: Spherical near-field (SNF) theory ：scattering matrix
SNF data
Measurand
Probe response coefficient
Spherical Near-Field Antenna Measurements Edited by J.E. Hansen
IEEE Std 1720-2012 IEEE Recommended Practice for Near-Field Antenna Measurements
rotation ceofficent
D.L. Meng, etc., Uncertainty Evaluation of NIM Spherical Near-field Antenna Calibration Facility, ACTA Metrologica SINICA, 38(6), 2017.
Measurement distance

## Slide 14: Comparison among Far field, Planar near field and Spherical near field
40GHz H面
26GHz E面
NIM: national institute of metrology, China.
CAST: China Academy of Space Technology
FF: far field method
SNF: Spherical near-field method
PNF: Planar near-field method
XP: cross polarization

## Slide 15: Compared between SNF ~ MoM (Method of Moment)
2.6 GHz E面
2.6 GHz H面
3.95 GHz E面
3.95 GHz H面

## Slide 16: Outlines
3
NIM Spherical near field facility
Uncertainty evaluation
4
Background & Problem
1
Methods for Calibrating antenna radiation pattern
2

## Slide 17: NIM SNF

## Slide 18: NIM SNF 400 MHz ~ 110 GHz
Enclosure：15m（L） × 7.6m（W） × 7.75m（H）
Rail length：15m
RF system
Anechoic chamber
Temperature ±1℃
Rail system, automatic cover system
Roll / Azimuth positioning
Laser interferometer
Alignment system for antennas and positioner
Software and documents
Sharing with extrapolation range
Notes: 频段范围-Dimension-rail(长度，直线度，定位精度)-激光测距仪-吸波材料自动覆盖系统-发射天线塔(双塔结构)-接收天线塔-测量系统（PNAx-）

## Slide 19: Precise Mounting & alignment
Mirror
Telescope
Mirror
Control Room
Anechoic chamber
Transit tower
Rotary joint
Text mixer
Positioner

## Slide 20: Example for mounting and alignment
75GHz～110GHz
400 MHz～490 MHz

## Slide 21: Positioner from Orbit/FR
AUC：
Max weight：250 kg
Max aperture：Ф2.5m
Orbit/FR AL-360-1P，option：
EN002(incremental encoder)
MM002 (metric threads in mounting hole interface)
TH002 (thru-hole for indoor operation)
CF001 (rear panel connector location)

## Slide 22: Reproducibility
26 GHz
40 GHz

## Slide 23: Outlines
4
Uncertainty evaluation
Background & Problem
1
Method for Calibrating antenna radiation pattern
2
NIM Spherical near field facility
3

## Slide 24: Fish-bone model for Radiation pattern with SNF

## Slide 25: IEEE Std 1720-2012: Principle of “self-comparison”
IEEE Std 1720-2012 IEEE Recommended Practice for Near-Field Antenna Measurements
Meng Donglin, etc., Uncertainty Evaluation of NIM’s Spherical Near-field Antenna Measurement Facility，ACTA METROLOGICA SINICA , 38(6), 2017
changes
changes
reference
Error
disturbances
Self-comparison：

## Slide 26: Measurement repeatability
26GHz
33GHz
40GHz

## Slide 27: Uncertainty evaluated with “self-comparison”
E-plane， 2.6 GHz
Changes from the tilting alignment of the AUC
E-plane， 3.95 GHz
Changes from the tilting alignment of the AUC

## Slide 28: Uncertainty from Multi-reflections
Notes: 三天线外推法天线测量的基础理论包括平面波散射矩阵，以及由此推导出的天线耦合方程和功率级数展开方程。实际测量时可依赖高精度导轨，测量由近及远一系列距离下收发天线之间的插入损耗，记为：𝑃(𝑑)=𝑃_R/𝑃_T ×1/𝐾 （1）。式中PR和PT分别是天线馈入功率和天线接收功率， 𝐾是收发天线馈电端口相连时的传输损耗，d是收发天线之间的距离。根据功率级数展开理论，P(d)d2可表示为：𝑃(𝑑) 𝑑^2=𝐴_1+𝐴_2/𝑑+𝐴_3/𝑑^2 +(𝐴_4/𝑑^3 +⋯)+多次反射项 （2），式中A1、A2、A3……An是功率级数展开系数。通过天线互耦抑制滤波算法滤除天线之间的多次反射和测量环境影响，基于有限阶多项式拟合算法忽略其中的高阶耦合项，从而得出有限个功率级数展开系数，如含3个系数的拟合结果为：𝑃(𝑑) 𝑑^2=𝐴_1+𝐴_2/𝑑+𝐴_3/𝑑^2 （3），当测量距离被外推至无限远处时，结合Friis传输公式，利用（3）中的拟合系数A1，可得收发天线在无限远处的增益乘积𝐺_𝑇 𝐺_𝑅。𝑑→∞，𝑝𝑑^2=𝐺_1 𝐺_2 (C/4𝜋𝑓)^2=𝐴_1 (4)。式中，GT和GR分别是发射天线和接收天线的增益，C是光速，f是频率。由式（4）可知：𝐺_T 𝐺_R=𝐴_1 (4𝜋𝑓/C)^2 （5）。以同样的方法，重复测量图2所示的三种收发天线组合的增益乘积，从而得到每个天线在无限远处的绝对增益[5]。

## Slide 29: Multi-reflections
E-plane， 3.3 GHz Multi-reflections
E-plane， 400 MHz, Multi-reflections

## Slide 30
Chart: Uncertainty of cross-polarization on the E-plane
2.6 GHz
Uncertainty from the cross-polarization
The SNF performance is depended on the probe’s x-performance
Large difference between -20 dB amd -30dB!

## Slide 31: The uncertainty budget of radiation pattern for AUT ETS 3106B (-30 dB to -20 dB below the boresight)
No. | Uncertainty source type | Source of uncertainty | Value |  |  | Probability distribution | Divisor
 |  |  | 400 MHz | 450 MHz | 490 MHz |  | 
1 | Probe | Radiation pattern (including polarization ratio) | 0.13 | 0.13 | 0.05 | Rectangular | 
2 |  | Multiple reflections between probe and AUT | 0.60 | 0.40 | 0.45 | Normal | 2
3 | Positioning | Probe misalignment | 0.18 | 0.20 | 0.10 | Rectangular | 
4 |  | AUT misalignment |  |  |  | Rectangular | 
5 |  | Scanning radius | 0.02 | 0.02 | 0.01 | Rectangular | 
6 |  | Blockage From positioner | 1.00 | 1.00 | 1.00 | Rectangular | 
7 | Data sampling | Measurement area truncation |  |  |  | Rectangular | 
8 |  | Data point spacing (aliasing) | 0.22 | 0.25 | 0.21 | Rectangular | 
9 | Room scattering | Room scattering | 2.50 | 2.00 | 1.25 | Normal | 2
10 |  | Scattering from the positioner |  |  |  |  | 
11 | Mismatch | Probe to source | - | - | - | - | -
12 |  | AUT to receiver | - | - | - | - | -
13 | RF system and accessories | Flexing cables/rotary joints | 0.60 | 0.50 | 0.25 | Rectangular | 
14 |  | Temperature effects |  |  |  | U shape | 
15 |  | Receiver errors |  |  |  | Rectangular | 
17 |  | Leakage and crosstalk | 0.05 | 0.35 | 0.13 | Rectangular | 
18 | Random | Random | - | - | - | - | 1
Combined standard uncertainty |  |  | 1.86 | 1.50 | 1.03 | Normal | 2
Expanded Uncertainty (k =2) |  |  | 3.8 | 3.0 | 2.1 |  | 2

## Slide 32: The uncertainty budget of radiation pattern for millitech SGH-10-RP1000 (-40 dB level below the boresight)
No. | Uncertainty source type | Source of uncertainty | Value |  |  | Probability distribution | Divisor
 |  |  | 75 GHz | 92 GHz | 110 GHz |  | 
1 | Probe | Radiation pattern (including polarization ratio) | 0.05 | 0.05 | 0.05 | Rectangular | 
2 |  | Multiple reflections between probe and AUT | 0.40 | 0.30 | 0.40 | Normal | 2
3 | Positioning | Probe misalignment | 0.80 | 0.70 | 0.80 | Rectangular | 
4 |  | AUT misalignment | 0.80 | 0.70 | 0.80 | Rectangular | 
5 |  | Scanning radius | 0.05 | 0.06 | 0.07 | Rectangular | 
6 |  | Blockage From positioner | 0.02 | 0.05 | 0.03 | Rectangular | 
7 | Data sampling | Measurement area truncation |  |  |  | Rectangular | 
8 |  | Data point spacing (aliasing) | 0.65 | 0.75 | 0.47 | Rectangular | 
9 | Room scattering | Room scattering | 0.02 | 0.02 | 0.02 | Normal | 2
10 |  | Scattering from the positioner |  |  |  |  | 
11 | Mismatch | Probe to source | - | - | - | - | -
12 |  | AUT to receiver | - | - | - | - | -
13 | RF system and accessories | Flexing cables/rotary joints | 0.20 | 0.30 | 0.20 | Rectangular | 
14 |  | Temperature effects |  |  |  | U shape | 
15 |  | Receiver errors |  |  |  | Rectangular | 
17 |  | Leakage and crosstalk | 0.10 | 0.35 | 0.20 | Rectangular | 
18 | Random | Random | - | - | - | - | 1
Combined standard uncertainty |  |  | 0.87 | 0.86 | 0.83 | Normal | 2
Expanded Uncertainty (k =2) |  |  | 1.8 | 1.8 | 1.8 |  | 2

## Slide 33: Conclusions
Introducing the role of metrology of antenna radiation pattern in remote sensing measurements
Demonstrate the accurate calibration of radiation pattern by PNF and SNF.
Detailed SNF facility in NIM is introduced; precise alignment system is key for good reproducibility.
The uncertainty of radiation pattern （amplitude) is introduced. Phase pattern will be researched in future.

## Slide 34: Thank you for your attention! Discussion?