## Slide 1
Collaborative Research of Applying JMA's GSICS Approach to FY-2G IR Channels w.r.t. Metop-A/IASI
Masaya Takahashi1, Yusuke Yogo1,
Qiang Guo2, Xiuqing Hu2, and Na Xu2
1Meteorological Satellite Center / Japan Meteorological Agency
2National Satellite Meteorological Center / China Meteorological Administration

## Slide 2
Backgrounds
In 2018, TB bias inconsistency in FY-2G IR channels vs. IASI-A between CMA-GSICS and GeoCAVS were shared by Qiang Guo with some GRWG members.
GeoCAVS: GSICS Algorithm + improved spatial collocation and parallax correction (Guo et al., 2018)
Tim Hewison suggested CMA/JMA collaborative research as we have a joint action on IR inter-calibration/comparison (A.GIR.2018.4c1).
JMA was pleased to apply JMA GSICS GEO-LEO-IR algorithm to FY-2G
  3-month validation of IR1-3 (10.8, 12.0, 6.95 μm) for MJJ 2017
Acknowledgements
Great thanks to Zhe Xu for data sharing and Peng Zhang for encouraging the collaboration.
Purpose of this study
To find root cause(s) of the inconsistency in CMA results
  Also helpful for JMA to validate our GSICS algorithm, and to find issues for future updates of the algorithm (e.g. regression method)

## Slide 3
JMA GSICS Algorithm for FY-2G/IASI-A
Based on JMA GEO-LEO-IR ATBD for MTSAT-1R/-2 and Himawari-8/-9
GEO-LEO collocation conditions
Gross error check is added in this study to easily remove outliers in regression
Updated IR2 SRF on 2016-03-22 (Guo and Feng 2017, RS) is used
Regression for inter-calibration coefficients (slope/offset): ordinary least squares
GEOradiance = slope LEOradiance + offset
 | Scene | IR1 (10.8 μm) | IR2 (12.0 μm) | IR3 (6.95 μm)
Obstime diff. | All | 5 minutes |  | 
Sat. zenith angle diff. | Clear | 0.01 |  | 
 | Cloudy | 0.03 |  | 0.01
Environment Uniformity [mW/m2/sr/cm-1] | Clear | 1.80 | 2.06 | 0.69
 | Cloudy | 3.59 | 4.11 | 0.69
Normality (Gaussian) | All | 2 |  | 1
Gross error [%] | All | 20 |  | 30
Clear/Cloudy scenes: roughly classified using 275 K of IR1 TB
(Y-axis) (X-axis)

## Slide 4
JMA GSICS Results
(Time window for stats: 15 days)
Same period of time as JMA validation
Guo (2018)
Same Y-axis range with CMA results
Results: FY-2G TB Bias@stdrad vs. Metop-A/IASI
CMA
CMA
IR1 (10.8)
IR2 (12.0)
IR3 (6.95)
TB biases among 3 methods
IR1 (10.8 μm)
GeoCAVS ≈ JMA ≈ CMA
IR2 (12.0 μm)
CMA << JMA ≈ GeoCAVS
IR3 (6.95 μm)
CMA ≈ JMA << GeoCAVS
Original IR2 SRF is used in CMA-GSICS

## Slide 5
JMA GSICS Results
(Time window for stats: 15 days)
IR2 TB Biases using Original SRF in JMA-GSICS
CMA
IR2 results: original SRF is used
FY-2G IR2 SRF
(Guo and Feng, 2017)
Consistent results between JMA/CMA GSICS
IR1 (10.8)
IR2 (12.0)
IR3 (6.95)

## Slide 6
Potential Reasons for IR3 Discrepancies btw GSICS/GeoCAVS
1) X/Y values for Regression
2) Impacts of outliers (gross error data)
3) Collocation Conditions: same GEO-LEO time diffs (5 min) in JMA-GSICS/GeoCAVS
4) Improved spatial collocation (variable weighting) and parallax correction in GeoCAVS
  Expected not to be the root cause of the IR3 discrepancy
JMA GSICS Results
(Time window for stats: 15 days)
CMA
Shown in next slides
Same Y-axis range with CMA results
IR1 (10.8)
IR2 (12.0)
IR3 (6.95)

## Slide 7
TB Biases at Standard Scene (247 K)
JMA (X-axis: IASI-A) : -1.70 K
JMA (X-axis: FY-2G) : -1.69 K
CMA (X-axis: FY-2G) : -1.69 K
Negligible difference at standard scene, but larger uncertainty at cold scenes
1) Impacts of X/Y values for Regression – IR3 (6.95 μm)
CMA results from CMA/NSMC GSICS website
JMA (X-axis: Metop-A/IASI) JMA (X-axis: FY-2G) CMA-GSICS (X-axis: FY-2G)
TB [K] TB [K]
FY-2G − IASI-A [K]
-6 -4 -2 0 2 4 6
200 220 240 260 280 200 220 240 260 280

## Slide 8
AAA
2) Impacts of Outliers on Regression: case for IR1 (10.8 μm)
 | w/o grserr check | w/ grserr check
@290 K | -0.156 (0.003) | -0.174 (0.003)
@250 K | -0.218 (0.012) | -0.482 (0.012)
@220 K | -0.332 (0.027) | -0.940 (0.027)
TB Biases [K]
(Std. err of TB)
TB [K]
FY-2G − IASI-A [K]
-6 -4 -2 0 2 4 6
200 220 240 260 280 300 320
TB [K]
200 220 240 260 280 300 320
FY-2G [mW/m2/sr/cm-1]
IASI-A [mW/m2/sr/cm-1]
FY-2G [mW/m2/sr/cm-1]
IASI-A [mW/m2/sr/cm-1]
w/o gross err check w/ gross err check
Red curve: estimated TB bias
from regression for 1-29 May 2017
Outliers: expected to be removed by referring quality flags in L1 data
Introduction of gross error check to GEO-LEO-IR: to be discussed later

## Slide 9
Summary
JMA GSICS algorithm was applied to FY-2G IR1-3 (10.8, 12.0, 6.95 μm) w.r.t. Metop-A/IASI to find the root cause of discrepancies between CMA-GSICS and GeoCAVS results.
Results/findings in this study (3-month validation for MJJ 2017)
Original IR2 SRF is used in CMA-GSICS
JMA-GSICS and CMA-GSICS: good agreement (0.1K) for IR1-IR3 under the condition of averaging during the 3-month period
GSICS and GeoCAVS: general agreement ( 0.3 K) for IR1 & IR2 under the condition of averaging during the 3-month period
GSICS Algorithm: negative bias in IR3 (> 1 K) compared with GeoCAVS
  Need for further investigation to find the root cause(s)
  Future tasks/collaborations: to be discussed in next slide

## Slide 10
Discussion
GSICS standard algorithm
ATBDs are publicly available if the products are submitted to GPPA
Products review in GPPA: ATBD, relevant documents, product file (file naming, metadata convention, values of variables)
Any requirements for additional reviews (e.g. code)?
Pros/cons of using gross-error check in GEO-LEO-IR
Robust/stable inter-calibration results: good for satellite data users
Losing chances for detecting cal. anomalies: no-good for satellite operators
Any requirements to generate two inter-calibration products?
Need for improving GEO-LEO-IR algorithm? by introducing:
Variable weighting, parallax correction suggested by Guo et al. (2018)
New spectral gap-filling method (discussed at June 2018 web meeting)
New regression method (discussed past web/annual meetings)

## Slide 11
Backup Slides

## Slide 12
JMA GSICS Algorithm for FY-2G/IASI-A
Based on JMA GEO-LEO-IR ATBD for MTSAT-1R/-2 and Himawari-8/-9
GEO-LEO collocation conditions
Observation time diff.
Satellite zenith angle diff. check:
Environment uniformity check:
Normality check
Gross error check (added in this study)
No need for Gap-Filling
Updated IR2 SRF on 2016-03-22 (Guo and Feng 2017, RS) is used
Regression for inter-calibration coefficients (slope/offset): ordinary least squares
GEOradiance = slope LEOradiance + offset
 | Scene | Thresholds
Obstime diff. | All | 5 minutes
Sat. zenith angle diff. (max_zen) | Clear | 0.01
 | Cloudy | IR1/2: 0.03 IR3: 0.01
Environment Uniformity [mW/m2/sr/cm-1] (max_STDV) | Clear | IR1: 1.80 IR2: 2.06 IR3: 0.69
 | Cloudy | IR1: 3.59 IR2: 4.11 IR3: 0.69
Normality (Gaussian) | All | IR1/2: 2 IR3: 1
Gross error [%] | All | IR1/2: 20 IR3: 30
/ |  100
Clear/Cloudy scenes: roughly classified using 275 K of IR1 TB
IR1: 10.8 μm
IR2: 12.0 μm
IR3: 6.95 μm
(Y-axis) (X-axis)

## Slide 13
TB Biases using NEW IR2 SRF TB Biases using ORIGINAL IR2 SRF
Time window for stats: 29 days (Re-Analysis Correction)
Impacts of New/Original IR2 SRFs on Inter-calibration
220 K 250 K 290 K StdRad

## Slide 14
TB Biases at Standard Scene (285 K)
0.45 K -0.64 K -0.29 K
CMA results from CMA/NSMC GSICS website
JMA using New SRF JMA using Original SRF CMA-GSICS
Impacts of New/Original IR2 SRF on Inter-calibration