## Slide 1
Monitoring SLSTR calibration using IASI: status and way forward
A. Burini I. Tomazic T. Hewison

## Slide 2: Introduction
Aim
  To verify if SLSTR IR is meeting the performance requirements:
    The absolute radiometric accuracy for data acquired by the IR channels shall be smaller than 0.2K, traceable to ITS-90.
    As a minimum, this requirement shall be met in the blackbody temperature range (~250K - ~300K)
Objectives
  To intercompare SLSTR with a stable and characterized reference sensor:
    IASI is the GSICS reference sensor due to its stability and chracterisation
    To follow GSICS methodology (Hewison, 2013)
    Based on EUMETSAT study by University of Leicester: Radiometric Intercomparison of AATSR and IASI, 2013, C. Whyte, D. Moore and J. Remedios
Collaboration with FIDUCEO project (H2020), Merchant et al.

## Slide 3: Input Dataset
SLSTR: @Se3-A and Sen3-B: SL_1_RBT___
IASI: @MetopA/B and C
  Level 1C
  SNOs time range
For atmospheric sounding
Spectral range: 645 to 2760 cm-1 (3.62 to 15.5 microns)
Maximum Optical Path Difference (OPD) : +/- 2 cm
Spatial resolution : pixel diameter of 12 km
Spectral resolution : 0.5 cm-1
Radiometric resolution : 0.2 to 0.4 K
Number of channels: 8461

## Slide 4: Methodology
Crossovers (SNO)
  SNO = simultaneous nadir overpasses
  based on orbital model prediction
  Each event last ~2 days with ~50 crossovers per SNO
Collocations (matchups)
  spatio/temporal/geometry
  Time: 5 min pixel time difference
  Spatial: within pixel
  Viewing angle: |Δsec(ϑ)/sec(ϑ)|<1% or |IASI SZA| < 20°
  IASI IFOV estimated as ellipse
Spectral convolution
  Convolve IASI radiance spectra with SLSTR SRF
  SLSTR SRF v2014
  S7 (3.7 µm) requires gap filling approach – incomplete IASI spectral coverage
Aggregate (average)
  SLSTR pixels in IASI FoV
Green – true collocation; blue – time of collocation + 3 min; red – time of collocation – 3 min
Limited to pole regions  temperature range ~200 K - 280 K

## Slide 5: S8 nadir: td5 min; |IASI sza|<20deg:no_signal filtered
Checkerboard pattern around ~200 K related to different response of det0 and det1
Removed all collocations containing SLSTR pixels flagged with no_signal  no tail
Indicates cal. problem for “valid” pixels: soon update of dynamic range (to 183 K) (S9 not affected)
ALL
Homogenous: (std<0.4 K within IASI IFOV)

## Slide 6: S8 nadir: td5 min; |IASI sza|<20deg:binned: ALL/homo: no_signal filtered
ALL
Homo
Jul16-Nov16
Jul16-Nov16

## Slide 7: S8: td5 min; |IASI sza|<20deg:binned: homo: no_signal filtered: detector0/1
S8: detector 0/1
S9: detector 0/1
Good agreement between detectors, but not the same!
Difference slightly more pronounced for S9
SLSTR - two IR radiometers!

## Slide 8
Objective:
  To solve the collocation problem in a general and fast way
  To create an infrastructure for the operational monitoring of IR sensor with IASI
  To extend the collocation criteria outside the polar region
  To ready for future missions
  To identify and test a “gap filling” method for S7 (extends beyond IASI range)
Re-engineering of the Prototype Processor

## Slide 9
Crossover
Based on Orbit Prediction
Identify matched only when there is an orbit intersection
Collocation
Space/temporal/angular based on STAMP (EUM)
Spectral Convolution
Aggregation of SLSTR Pixels in IASI FOV
Generation of potential candidates couple for the analysis
Aggregation
Convolution of IASI radiance spectra with SLSTR SRF
Processing Schema
Data Access Interface
Based on database and catalogues
Collocation
Minimization of distortion
Based on Proj4 + GIS libraries
Spectral Convolution
To use the GSICS format for SRF
Identification of matched data outside the polar regions
Stand Alone Service
Aggregation
IASI FOV modelling – Generic Modelling of FOV projected on Earth surface
Connection to Database for storing results

## Slide 10
Data Access as a service
Cataloguing of data (testing of ElasticSearch or classical DB solutions)
Daemon tool to perform cross query:
  Identification of Potential Candidates
  Overlap Criteria can be customized (percentage of overlapping polygons)
Tool Completely de-coupled from the processing:
  To speed up the pre-collocation process
  It can be re-used for other processes

## Slide 11
A Generic Collocation Approach
To make re-use of well tested and broadly accepted libraries for geographic projection, aiming to minimize the distortion of data
Collocation between sounding sensors and Imaging sensor is then automatically achieved (same reference system) = DATA CUBE
The collocation is done at “granule” level (or at scenario level ) in order to be easily parallelized over several nodes. (it can still be applied in the GEO disk)
Projection in a common reference grid.
Orthographic projection
Oblique Mercator
Transversal Mercator

## Slide 12: Summary
Prototype for monitoring SLSTR calibration with IASI as reference has been developed. Thanks to it the following considerations can be made:
  Stable calibration within requirements (220-280 K)
    Near zero nadir view bias (<0.1 K) in S8 and S9 (220 K – 280 K)
    Oblique view looks ok – to be confirmed with QSNO
  Cold temperature bias (~>0.2 K)
    Nonlinearity? - BB 250-300 (also observed with the “Tandem” analysis)
    S8 cold temp. calibration issue ~200 K (dynamic range, “no_signal”)
    Small difference in dynamic range for two detectors (will be lowered to ~183 K)
  Different detectors response:
    S8 – OK, S9 small offset between detectors (~0.08 K)
    SLSTR IR  2 radiometers
Preparation for a more generic software suite and generic functions has started. Planned to be concluded in few months.