## Slide 1: S3 SLSTR TIR Calibration
Dave Smith – SLSTR Calibration Scientist
Ed Polehampton, Mireya Etxaluze
RAL Space, Science and Technologies Facilities Council, United Kingdom

## Slide 2: Sea and Land Surface Temperature Radiometer (SLSTR)
2021 – Sentinel 3C
  Spectral Calibration in progress
  Instrument Calibration Spring 2019
2023 – Sentinel-3D
…
2016 – Sentinel 3A
Launched 16-Feb-2016 
2018 – Sentinel 3B
Launched 25-Apr-2018 

## Slide 3: ATSR Series
1991-2000 ATSR-1
1995-2008 ATSR-2
2002-2012- AATSR

## Slide 4: SLSTR instrument
Nadir swath >74° (1400km swath)
Dual view swath 49° (750 km)
Two telescopes 110 mm / 800mm focal length
Spectral bands TIR : 3.74µm, 10.85µm, 12µm
SWIR : 1.38µm, 1.61µm, 2.25 µm
VIS: 555nm, 659nm, 859nm
Spatial Resolution 1km at nadir for TIR, 0.5km for VIS/SWIR
Radiometric quality NEΔT 30 mK (LWIR) – 50mK (MWIR)
SNR 20 for VIS - SWIR
Radiometric accuracy 0.2K for IR channels
2% for Solar channels relative to Sun

## Slide 5: On-Board Calibration systems
Thermal InfraRed
Blackbodies
VIS-SWIR Channels
VISCAL
Effective e >0.998
T non-uniformity < 0.02 K
T Abs. Accuracy 0.07 K
T stability < 0.3 mK/s
8 PRT sensors + 32 Thermistors
Zenith diffuser + relay mirrors
Uncertainty <2%

## Slide 6: The Goal
To ensure the interoperability of satellite datasets it is a requirement for their measurements to be calibrated against standards that are traceable to SI units
For temperature this is the International Temperature Scale of 1990
For IR instruments such as SLSTR the traceability is achieved via internal BB sources
S-PRT
Fixed Point Cells
Blackbody Source
Instrument

## Slide 7: SLSTR L1 Processing
Processing specification defined by
  ATBD -> DPM
  L0 and L1 Product Specifications
Each spectral band (5 thermal bands) and detector element (2x2) for each for each earth view (separate for nadir and oblique) has unique set of calibration calibration coefficients
  = 40 for IR channels alone
Contained in Satellite Characterisation and Calibration Database Document (S-CCDB)
  Configuration controlled by MPC

## Slide 8: Calibration Model
We obtain calibration coefficients via reference to known calibration sources
Typically detector counts will be some function of the earth scene radiance
which reduces to
Both gain and offset must be stable during calibration interval
We invert this to get scene radiance as a function of detector counts
(assuming linear function)
Uncertainty in scene radiance

## Slide 9: SLSTR TIR Calibration Effects Tree

## Slide 10: Primary Sources of Uncertainty in TIR calibration
Black-Body Temperatures
  PRT calibration at subsystem level – traced to SPRT (ITS-90) – NPL/NIST
  Blackbody gradients, thermal analysis - RAL
Black-Body Cavity Emissivity
  Spectral Reflectance of Black Coating – NIST/NPL
  Cavity Model – STEEP323 or SMART3D (ABSL model)
Spectral Response
  FPA measurements – RAL reports [S3-RP-RAL-SL-102 (S3A), S3-RP-RAL-SL-114 (S3B)]
Non-Linearity
  Instrument level calibration tests – RAL reports
Detector Noise
  Instrument level calibration tests, on-board BB sources

## Slide 11: S3B TIR Uncertainty Budget – Jun-2018
Uncertainties vs. Temperature from pre-launch are included in L1 products

## Slide 12: Uncertainties in L1 products
S8 BT
S8 NEDT
(Random Effects)
S8 Calibration Uncertainty (Correlated Effects)
RGB False Colour
S3A_SL_1_RBT____20171129T014030_20171129T014330_20171129T032538_0179_025_074_3059_SVL_O_NR_002.SEN3
L1 Products contain information to generate per pixel uncertainty estimates
Noise estimates (random) from on-board sources
Uncertainty in calibrated (correlated) is derived from pre-launch calibration
User tool under development by RAL to generate uncertainty ‘images’

## Slide 13: Pre-Launch Calibration - SLSTR-A
Nadir View
Oblique View
Plots show the differences between the BT measured by SLSTR and the BT derived from the BB thermometry
Uncertainties are combined uncertainty estimates – k=1
Only S7-S9 shown for clarity

## Slide 14: Pre-Launch Calibration - SLSTR-B
Nadir View
Oblique View
Plots show the differences between the BT measured by SLSTR and the BT derived from the BB thermometry
Uncertainties are combined uncertainty estimates – k=1
Only S7-S9 shown for clarity

## Slide 15: Why the differences?
Non-Blackness of optical stops (i.e. ɛ < 0.9) causing non-uniform thermal background
Measurements by PTB confirm issue with black coatings
Hence modification to stop coatings
Temperature gradients in flight BBs
Thermal modelling shows asymmetry of baseplate temperatures
Analysis using monte-carlo model (STEEP-323) suggests that non-uniformity does not affect BB radiances

## Slide 16: Post Launch Calibration – The Challenge
Field radiometers validate the L2 product (SST, LST)
  Dominated by uncertainties in atmospheric correction, surface emissivity, temporal effects, retrieval algorithm
There are no stable reference sites!
  TIR radiances are influenced by temperature which does vary over short time scales
  Atmosphere contributes to the measured radiances – strong spectral sensitivity – even in window bands
Post Launch Calibration is therefore limited to:
  Monitoring of instrument temperatures – in particular detectors, optics,
    BB-sources since these contribute to the uncertainty budget.
  Comparisons with other sensors – i.e. IASI
    Limited to SNOs
  Tandem phase analysis

## Slide 17: Detector Temperatures
IR detector temperatures maintained between 84 K and 87 K
Periodic de-contamination is needed to remove water ice from cold surfaces
SLSTR-A FPA Cooler set-point increased by 1K in July to increase running time between decontaminations
SLSTR-A
SLSTR-B
VIS Detectors
VIS Detectors
IR Detectors
IR Detectors

## Slide 18: Black-Body Performance
Blackbody temperatures have a seasonal cycle on top of the daily/orbital temperature cycles. Heated BB remains below 305K limit necessary for S7 calibration.
Temperature gradients consistent with pre-launch values for SLSTR-A and B
Note higher gradient in heated BB for SLSTR-A
SLSTR-A
SLSTR-B
+YBB
-YBB

## Slide 19: IR Channels Gains
All detectors are performing as at launch within expected ranges
Gains and offsets show slow drift with detector temperature as expected.
SLSTR-A
SLSTR-B
S7
S8
S9
F1
F2

## Slide 20: IR Channels Noise
Radiometric noise levels for the TIR channels have remained stable throughout at pre-launch values.
NEDT for the S8 and S9 channels are below 20mK with no indication of degradation.
S3B F1 shows periodic increases in noise – possibly due to motional chopping
SLSTR-A
SLSTR-B
S7
S8
S9
F1
F2

## Slide 21: IASI-SLSTR Co-Locations from Eumetsat SLSTR-A/B vs IASI-A
ALL SNO events
Simultaneous Nadir Observations (SNO) performed by Eumetsat following GSICS methodology based on the approach developed for AATSR/IASI comparisons.
  IASI data are spectrally averaged over SLSTR bands
  SLSTR data are spatially aggregated over IASI footprint
Results shown for SNO event (20-22.06.2018) during the S3 tandem phase
Analysis for both S3A and S3B are performed separately for the two detectors of the two nadir view channels S8 and S9
Comparisons are shown for homogenous scenes (∆T <0.1K)
SLSTR-A vs IASI-A
Nadir: 12 µm (S9)
SLSTR-B vs IASI-A
Nadir: 10.8 µm (S8)
SLSTR-A vs IASI-A
Nadir: 10.8 µm (S8)
SLSTR-B vs IASI-A
Nadir: 12 µm (S9)
Courtesy I.Tomazic (Eumetsat)
SLSTR-A vs IASI-A
SLSTR-B vs IASI-A

## Slide 22: Tandem Phase Analysis
Sentinel-3A flying on Same Track 30s behind Sentinel-3B
Sentinel-3A
Sentinel-3B

## Slide 23: Tandem Phase Analysis
Note comparisons are affected by spatial variations – e.g. clouds, surface features
S3 (870nm) Differences
S9 (12um) Differences

## Slide 24: TIR Channel Radiometric Comparisons
Comparisons of S3B vs S3A have been performed for selected days during tandem phase using L1 products remapped to 0.05degree cells.
  Comparisons are for corresponding 0.05degree cells
Averages are filtered to use spatially uniform cells with standard deviations <0.1K and BTs within valid range.
  Invalid BTs are not included in averages
Summary plots have been averaged over 5K intervals

## Slide 25: SLSTR-A S8 – Night Time - Nadir
Data for 26-August-2018
Notes: L1 data are averaged in 0.05degree cells as for VIS/SWIR channels

## Slide 26: SLSTR-B S8 – Night Time - Nadir
Data for 26-August-2018
Notes: L1 data are averaged in 0.05degree cells as for VIS/SWIR channels

## Slide 27: S3B-S3A Tandem Comparisons
Level-3 data for 26th August 2018
BT Differences S3B - S3A
S7
S8
S9

## Slide 28: S3B-S3A Comparisons – Nadir View
Low temperature differences possibly due to residual non-linearity
Unfiltered
±10% Outliers removed

## Slide 29: S3B-S3A Comparisons – Oblique View
Error bars are standard deviation of match-ups
Unfiltered
±10% Outliers removed
Notes: Add comment on stray light

## Slide 30: Stray Light Correction
The radiometric calibration model depends on the thermal background being constant around the scan
What if the thermal background is different between the earth view and the hot and cold blackbody sources?
Recalling the calibration model
      where
We make use of the +0 term in the BB radiance model to get

## Slide 31: Potential Impact of Correction on S3B-S3A Comparisons
As measured
With Correction
The model shows that the differences between SLSTR-A and B are consistent with the pre-launch calibration errors and can be corrected with a reduced version of the stray-light model developed for the IPF-P.
Notes: Add comment on stray light

## Slide 32: Conclusions
Sentinel-3A/B Tandem phase has provided a unique opportunity to cross-compare twin instruments in near ideal conditions
Preliminary analysis shows:
  Good agreement between S3B and S3A in nadir view
  Differences in oblque view consistent with pre-launch calibration
  Both nadir and oblique show residual non-linearity error for TE below 220K (note design range 220K – 330K)
SNO comparisons are so-far limited to match-ups with IASI-A
  Limited to high latitude observations – TE < 280K
  Cross calibration with other sensors would be useful – VIIRS, AIRS…
Study with NPL is in progress to document and publish current status of the traceability of SLSTR TIR channels

## Slide 33: Spares

## Slide 34: IR Calibration Model – Noise Estimates
Noise estimates (solid red line) were predicted using instrument noise model and the expected performance of FM Detector + EM Electronics + margins for EOL performance.
Measured values (points) are from pre-launch calibration representative of ‘cold’ start of mission

## Slide 35: Modelling Calibration Error
A small team was set up to investigate possible causes of the discrepancies in the calibration results
  I.e. what are the additional terms in the calibration model.
An analysis suggested a non-uniform background radiance which varied around the scan.
  BB radiance errors
  Non-Black coating on optical baffles
The effects were modelled and a correction was tested with the data.
  Model is has several free parameters and is an approximation
Early inter-comparsions with IASI performed by EUMETSAT suggest that on-orbit stray-light error correction is not necessary.
Nadir View

## Slide 36: Detector Temperatures
IR detector temperatures maintained between 84 K and 87 K
Periodic de-contamination is needed to remove water ice from cold surfaces
SLSTR-A FPA Cooler set-point increased by 1K in July to increase running time between decontaminations
SLSTR-A
SLSTR-B
VIS Detectors
VIS Detectors
IR Detectors
IR Detectors

## Slide 37: Black-Body Performance
Blackbody temperatures have a seasonal cycle on top of the daily/orbital temperature cycles. Heated BB remains below 305K limit necessary for S7 calibration.
Temperature gradients consistent with pre-launch values for SLSTR-A and B
Note higher gradient in heated BB for SLSTR-A
SLSTR-A
SLSTR-B
+YBB
-YBB

## Slide 38: IR Channels Gains
All detectors are performing as at launch within expected ranges
Gains and offsets show slow drift with detector temperature as expected.
SLSTR-A
SLSTR-B
S7
S8
S9
F1
F2

## Slide 39: IR Channels Noise
Radiometric noise levels for the TIR channels have remained stable throughout at pre-launch values.
NEDT for the S8 and S9 channels are below 20mK with no indication of degradation.
S3B F1 shows periodic increases in noise – possibly due to motional chopping
SLSTR-A
SLSTR-B
S7
S8
S9
F1
F2

## Slide 40: SLSTR-A S8 – Night Time - Oblique
Data for 26-August-2018
Notes: L1 data are averaged in 0.05degree cells as for VIS/SWIR channels

## Slide 41: SLSTR-B S8 – Night Time - Oblique
Data for 26-August-2018
Notes: L1 data are averaged in 0.05degree cells as for VIS/SWIR channels

## Slide 42: S3B-S3A Tandem Comparisons Oblique
Level-3 data for 26th August 2018
BT Differences S3B - S3A
S7
S8
S9