## Slide 1: Application and evaluation of New NOAA gap filling for CMA IR hyperspectral instruments
Hanlie XU, Na XU, Xiuqing HU
CMA
2019-03-07

## Slide 2: Outline
Spectral Coverage Comparison
Application testing based on SNO intercalibration
  MERSI-CrIS
  MERSI-HIRAS
Summary

## Slide 3: Instrument—MERSI-II
MERSI-II Improvements:
Cover all bands in FY-3A/B/C MERSI and VIRR
Five more IR bands
Cirrus cloud band 1.38μm
Water vapor bands In NIR and 7.2μm
Two IR split windows with 250m spatial resolution
Higher accuracy from onboard calibration
Lunar Calibration capability
Band | FY-3D MERSI-II | FY-3A/B/C MERSI
1 | 0.470 | 0.470
2 | 0.550 | 0.550
3 | 0.650 | 0.650
4 | 0.865 | 0.865
5 | 1.38 | 11.25
6 | 1.64 | 1.640
7 | 2.13 | 2.130
8 | 0.412 | 0.412
9 | 0.443 | 0.443
10 | 0.490 | 0.490
11 | 0.555 | 0.520
12 | 0.670 | 0.565
13 | 0.709 | 0.650
14 | 0.746 | 0.685
15 | 0.865 | 0.765
16 | 0.905 | 0.865
17 | 0.936 | 0.905
18 | 0.940 | 0.940
19 | 1.03 | 0.980
20 | 3.8 | 1.030
21 | 4.05 | 
22 | 7.2 | 
23 | 8.550 | 
24 | 10.8 | 
25 | 12.0 | 
250 m
1000 m

## Slide 4: Instruments Spectral Coverage
CH25
CH24
CH23
CH22
CH21
CH20
FY-3D/MERSI |  | FY-3D/HIRAS | NPP/CrIS | METOP/IASI
Band number | Wavelength Center | Spectral channel coverage (cm-1) |  | 
CH25 | 12μm (833cm-1) | 650-1135 | 650-1095 | 645-1210
CH24 | 10.8μm (926cm-1) |  |  | 
CH23 | 8.55μm (1169cm-1) | Gap | Gap | 
CH22 | 7.2μm (1389cm-1) | 1210-1750 | 1210-1750 | 1210-2000
CH21 | 4.05μm (2469cm-1) | 2155-2550 (A little bit gap) | 2155-2550 (A little bit gap) | 2000-2760 (A little bit gap)
CH20 | 3.8μm (2631cm-1) | Gap | Gap | 
MERSI SRF & IASI、HIRAS spectral coverage
Spectral Gap of HIRAS

## Slide 5
From Hui Xu
huixu@umd.edu
Gap filling Coefficients
Usage : There are 3 datasets inside the H5 file
The use of the H5 file
t observations
1158 gap radiances
t observations
2211 measured radiances
2211 coefficients
1158 gap channels
1158 gap channels
CrIS-FS
measured spectra
CrIS-FS
gap spectra
Methodology:Principal Component Regression (PCR) based spectral gap prediction
P0：is gap-filling coefficient matrix with a dimension of 1158(col) ×2211(row)
C0：is gap-filling constant vector with a length of 1158
GAP_NUM: has 3 values which are 183, 647 and 328 representing the number of gap channels (totally are 1158=183+647+328) in LW[1095.625~1209.375], MW [1750.625~2154.375] and SW [2550.625~2755] spectral regions respectively.

## Slide 6: CrIS Radiance Spectra after spectral filling
Application Testing: MERSI-CrIS
Red ：Orignal Spectrum
Blue：Filling Spectrum

## Slide 7: SNO Intercalibration Regression based on IASI and CRIS for MERSI Band 8.55μm
S
Before
After
Bias@250K=-0.39K
Bias@250K=-0.54K
CrIS Spectral Gap
Application Testing: MERSI-CrIS
MERSI SRF
SNO track selection
60°N~80°N
Radiance Bias
TBB Bias

## Slide 8
TBB Bias
Double difference (CrIS&IASI-B)
CRIS(filling)
IASI-B
8.55μm
10.8μm
Application Testing: MERSI-CrIS
8.55μm
8.55μm
FY4A/AGRI 7.1μm

## Slide 9
20180601-0820
Warm Scene
Blue：Orignal Spectra
Red ：Gap filling Spectra
Application Testing: MERSI-HIRAS
HIRAS Radiance Spectra after gapfilling
Cold Scene

## Slide 10: Details spectra
Warm Scene
Cold Scene
Application Testing: MERSI-HIRAS
LW and MW
MW and SW

## Slide 11
CrIS Gap
HIRAS Gap
SNO Intercalibration Regression based on CrIS and HIRAS for MERSI Band 8.55μm
MERSI SRF
Application Testing: MERSI-HIRAS

## Slide 12
Summary and Discussion
Testing results of CrIS Spectral filling method on FY-3D MERSI 8.55μm band and HIRAS show the new spectra after gap filling is reliable, and can be used as reference for inter-calibration for the broadband channel.
Why at the CrIS spectral gap area, the CrIS is cooler than IASI?
Likun Wang,et al., GSICS annual meeting,2016
FY4A/AGRI 7.1μm
FY3D/MERSI 10.8μm

## Slide 13
Thank for your attention