## Slide 1
1
S-NPP, N-20, and N-21 VIIRS Reflective Solar Bands
On-orbit Radiometric Calibration Performance
VIIRS Characterization Support Team (VCST), NASA GSFC
(presented by Ning Lei)
March 12, 2024
GSICS Annual Meeting, March 11-15 Darmstadt Germany

## Slide 2
2
Outline
VIIRS reflective solar band (RSB) on-orbit radiometric calibration
events and methodology
2. RSB on-orbit calibration performance update
3. Summary

## Slide 3
3
Major events/updates since GSICS 2023
Spacecraft Command and Data processor reset (July 26, 2023)
Updated the HRTA model parameter values
S-NPP
Spacecraft Command processor resets (Sep. 29, 2023, Feb. 4, 2024)
VIIRS Command processor reset to resolve a quality flag bit issue (Feb. 20, 2024)
Updated the HRTA model parameter values
Implement de-striping algorithms for VisNIR bands
N-20
Refined the screen transmittance functions, using calibration data collected
on both the yaw maneuver and regular orbits
Second mid-mission outgassing to restore SWIR and thermal emissive bands to at-launch performance levels
SDSM operation frequency changed from one orbit per day to one orbit per week
N-21

## Slide 4
4
VIIRS RSB on-orbit calibration
Space view
SD screen
SDSM screen
sunlight
SDSM
RTA
Solar Diffuser (SD)
Solar Diffuser (SD): a calibration source; its BRDF change (H-factor) measured by the SD stability monitor (SDSM)
on-orbit calibrated
Spectral radiance

## Slide 5
5
Calculating F(SD) needs HRTA
Solar azimuth angle
S-NPP and N-20
N-21
Use a sliding window (in time) approach, fit F(HSDSM) to
F(Moon) to find
(Note: once more lunar F-factors are available, we’ll use the methodology for SNPP and N20)

## Slide 6
6
Performance of S-NPP SD and SDSM
H-factors dropped by about 1% due to the Feb. 24, 2019 event, faster at a shorter wavelength
Recently (~last 1.5 years), H-factors trend upwards, also occurring for N20 and MODIS SDs
SDSM detector gains trend normally; largest decrease for SDSM detector 8 (longest wavelength)—lattice damage from solar bombardment
Feb. 24, 2019 event

## Slide 7
S-NPP RSB F-factors
Solid lines: 1/F(SD); circles: 1/F(moon)
F decreases with time, because of telescope mirror surface tungsten oxide contamination
Over the last 3-4 months, F(SD) diverges from F(moon): need an investigation

## Slide 8
H-factors decrease at slower rates than S-NPP H-factors
SDSM gains decrease similarly to S-NPP
Performance of N-20 SD and SDSM
Notes: 8
SDSM_Hplot_V4aeroX.20150629

## Slide 9
N-20 RSB F-factors
F-factors are very stable over mission
Over the last 3-4 months, F(SD) diverges from F(moon): need an investigation
Solid lines: 1/F(SD); circles: 1/F(moon)

## Slide 10
H-factors decrease faster than N-20, but similarly to S-NPP H-factors
SDSM gains decrease similarly to S-NPP and N-21
Performance of N-21 SD and SDSM
Notes: 10
SDSM_Hplot_V4aeroX.20150629

## Slide 11
Performance of N-21 RSBs: F-factors
The VisNIR band SD gains (1/F; solid lines)) are nearly flat and closely follow the respective lunar gains (circles)
2nd mid-mission outgassing return the SWIR band gains to their original values
2nd mid-mission outgassing

## Slide 12
SNRs
All SNRs satisfy requirements, except N-20 I3 band detector 29 (noisy detector)

## Slide 13
13
S-NPP and N-20 H-factor SD positional dependence
SD BRDF change factor (H-factor) can depend on SD position
detector array
model parameters

## Slide 14
M1 striping not seen (C2.0)
M1 striping (C1.1)
2019229
H-factor SD positional dependence implemented is still good
S-NPP TOA reflectance images
2024019
2019229
M1: no striping seen (C2.0)

## Slide 15: Summary
S-NPP, N-20, and N-21 VIIRS RSBs perform normally with SNRs satisfying specifications (except N-20 I3 detector 29) and are expected to be above specifications for the foreseeable future
N-20 and N-21 RSB VisNIR band gains are relatively stable with time whereas S-NPP RSB gains depart from prior trend for some bands since July 2022
S-NPP and N-21 SDs degrade at similar rates, faster than N-20 SD
All three VIIRS instruments SDSM detector gains decrease with similar rates at the same wavelength
De-striping algorithm applied to S-NPP RSBs still works; de-striping algorithms applied to N-20 VisNIR bands recently
Newly derived N-21 VIIRS screen functions significantly improve the accuracy of the retrieved H- and F-factors