## Slide 1: OMI (and TROPOMI) Solar Activity Time Series
Matthew DeLand1, Sergey Marchenko1
1Science Systems and Applications, Inc. (SSAI)
GSICS UV/VIS/NIR Topical Meeting
Virtual 30 May 2024
Supported by NASA Grant 80NSSC21K1946

## Slide 2: Aura Ozone Monitoring Instrument (OMI)
Launched 15 July 2004; More than 105,000 orbits so far!
Spectral coverage: 265-500 nm
  UV1 band: 265-309 nm, Δλ = 0.63 nm
  UV2 band: 309-380 nm, Δλ = 0.42 nm
  VIS band: 380-504 nm, Δλ = 0.63 nm
Nadir-viewing pushbroom spectrometer with wide field of view  30 (UV1) or 60 (UV2, VIS) cross-track views per image
Solar measurements use diffusers (quartz volume, aluminum reflective) in approximately daily [15 orbits] and monthly cadence
Optical design (spectral shift between cross-track views) gives highly oversampled solar irradiance spectrum

## Slide 3: OMI Advantages for SSI Measurements
Low overall degradation is easier to correct to high precision
  10% at 270 nm, 4% at 400 nm during mission lifetime
Average many spectra for excellent signal-to-noise performance
  ~250-500 spectra per daily sequence
Spectral resolution helps characterization of solar activity in key absorption features (Mg II, Ca II, molecular bands)
  Δλ = 0.4-0.6 nm
Long-term wavelength stability
  0.02-0.03 nm drift during mission lifetime

## Slide 4
Stability / year | 0.05% (λ < 400 nm), 0.01% (λ > 400 nm)
Relative precision | 0.01%
O. Coddington et al., 2016, BAMS, 97, 1265
Daily OMI Solar Cycle 24 SSIs (2006-present, V3 [Marchenko et al., 2019]):
* Comparisons between adjacent solar minima, i.e., relative change
Stability / year * | ≤ 0.02% (λ < 400 nm), ≤ 0.01% (λ > 400 nm)
Relative precision | ~ 0.1% (λ < 400 nm), ~0.02-0.1% (λ > 400 nm)
SSI Requirements for Climate Studies
Long-term (decades), uninterrupted SSI records
Preferably ~daily SSI observations
We have developed alternative L1 degradation corrections to try to meet these goals
Notes: A tall order indeed!
the numbers in the top half represent the NOAA Climate Data Record requirements that TSIS (next slide) is hoping to meet

## Slide 5
faculae
sunspots
1%
OMI SSI Time Series (Cycle 24 band averages)
Marchenko et al. [2019]
Notes: - All the cycle-24 history is in 1%
Mostly, the ‘noise’ is a real solar variability
OMI, besides SORCE, is the only instrument capable of delivering the long-term SSIs of desired quality

## Slide 6: OMI [V6] Comparison to TSIS-1 SIM [V12]
Use narrow spectral bands (Δλ = 5-10 nm)
Agreement between concurrent measurements is within long-term stability (0.1-0.2%)

## Slide 7
Solar Activity: Mg II Line Cycle 24
0.5%
OMI Mg II index vs. SORCE SOLSTICE Mg II index

## Slide 8
Solar Activity: Mg II Line Cycles 24-25
1.5%
OMI Mg II index vs. Bremen Mg II index
OMI Mg II index vs. GOES-16 Mg II index

## Slide 9
OMI Ca II K index vs. SFO Ca II K products
Solar Activity: Ca II Lines Cycles 24-25

## Slide 10
No adjustments made for slightly different spectral resolution between OMI and TROPOMI
OMI and TROPOMI Ca II index products agree to ~0.03% on average
Solar Activity: Ca II Lines Short-term

## Slide 11
Use additional TROPOMI bands (658-782 nm, 2300-2388 nm) to construct indexes using Fe I, CO lines
Identify short-term signals down to ~0.02% peak-to-peak response
Examine spectral dependence (e.g. quiet activity at 700-800 nm)
Solar Activity: Other Indexes

## Slide 12: Future Plans
OMI operations have been approved through mid-2025
  Aura orbit drift is causing periodic goniometric calibration issues for solar irradiance measurements
  First episode (October-December 2023) has been treated in our SSI product
TROPOMI operations continue
  Ca II index product is being archived at LISIRD site
  Artifacts in released irradiance data (particularly for λ < 300 nm) exceed desired threshold for long-term use
  We are working with KNMI team to understand these features
  Also hope to get access to uncorrected irradiance data for possible development of alternative degradation correction