## Slide 1: LIME: Lunar Irradiance Model of ESA Sarah Taylor (NPL), Stefan Adriaensen (VITO), Carlos Toledano (UVa), Alberto Berjón (UVa), Africa Barreto (UVa), Emma Woolliams (NPL), Maria Garcia-Miranda (NPL), Marc Bouvet (ESA)
Welcome to the National Physical Laboratory

## Slide 2: LIME: Lunar Irradiance Model of ESA
Model Overview
Instrument description and Calibration
Derivation of TOA Irradiance
Model Regression
Initial Model Comparisons
Project management
Instrument calibration and uncertainty analysis
Deployment of Photometer
Building database of lunar irradiance Measurements
Lunar Model Development
Model comparisons

## Slide 3: LIME Overview
Derived using SI-traceable ground-based measurements acquired with CIMEL 318-TP9 photometer from high altitude location at Teide Peak and Izaña Atmospheric Observatory in Tenerife
Characterization and calibration at NPL and University of Valladolid
Rigorous uncertainty analysis for calibration and individual observations
Some modifications on original ROLO model and uncertainty estimated from Monte Carlo simulations.

## Slide 4
CIMEL 318-TP9

## Slide 5: Linearity (NPL)

## Slide 6: Temperature Sensitivity (UVa)
Pico Teide | Jan | Feb | Mar | Apr | May | Jun | Jul | Aug | Sep | Oct | Nov | Dec
Percentile 98 [°C] | 7.7 | 8.0 | 8.7 | 11.7 | 14.1 | 16.6 | 18.7 | 18.8 | 15.1 | 13.0 | 9.9 | 7.5
Percentile 2 [°C] | -9.5 | -12.0 | -8.1 | -5.7 | -2.7 | 1.8 | 5.0 | 5.0 | 1.9 | -3.5 | -5.9 | -8.2
Notes: The temperatures at the observation sites clearly varies significantly,
As the instrument is not thermally stabilised, the thermal sensitivity of the instrument must be determined. This is particularly important for 1020 nm, where Si is particularly temperature sensitive but is also necessary to determine for all spectral channels.
Characterisation allows determination of temperature correction coefficients which are applied to all observations and the calibration which is performed at known temperature.
The standard instrument operating temperature range of the photometer provided by the CIMEL Electronique is -30 °C to +40 °C (CIMEL, 2015).
The temperature coefficients for the photometer are determined in a temperature chamber CLIMATS TM. It consists of a stainless steel cabinet with a gridding where the photometer can be positioned, and the right side has an 80 mm Ø hole. The photometer is aligned with an integrating sphere (Labsphere 10’’ diameter), illuminated with a 100 W lamp, and powered by a stabilized power supply (Agilent E3634A). The stabilisation time for the integrating sphere is 20 minutes.
The temperature of the chamber is varied from 50 °C to -40 °C during a period of 4 hours, which is a rate of change of about 0.3 °C/min. The temperature ramp is started when instrument temperature reaches 40 °C.
The irradiance, 𝐸_𝑇 measured in a temperature chamber at different temperatures 𝑇 [°C], while the detector is illuminated with a stable light source, is fitted to the following model:
Where 𝑇_ref is the reference temperature (25 °C) and 𝐸_𝑐 is the irradiance of the source measured at the reference temperature.

## Slide 7: Irradiance Responsivity
Spectral Channel | Calibration Coefficient (MOON gain) | Standard uncertainty
440 nm Si | 5.759 × 10-10 | 0.97%
500 nm Si | 4.481 × 10-10 | 0.96%
675 nm Si | 3.205 × 10-10 | 0.92%
870 nm Si | 2.547 × 10-10 | 0.91%
937 nm Si | 2.431 × 10-10 | 0.97%
1020 nm Si | 2.735 × 10-10 | 1.05%
1020 nm InGaAs | 2.119 × 10-10 | 1.01%
1640 nm InGaAs | 4.893 × 10-11 | 1.06%

## Slide 8: CIMEL Stability by Solar Langley Calibration
WLN​ | 23/06/2018​ | 25/06/2020​ | Difference(%)​
1020​ | 628596.4​ | 631346​ | 0.44​
1640​ | 1114638.9​ | 1114576​ | -0.01​
870​ | 882649.8​ | 885586​ | 0.33​
675​ | 1111487.4​ | 1114076​ | 0.23​
440​ | 760622.6​ | 761067​ | 0.06​
500​ | 1026944.7​ | 1029153​ | 0.21​
1020i​ | 805945.7​ | 808002​ | 0.25​
935​ | 827605.4​ | 814508​ | -1.60​
380​ | 129951​ | 126758​ | -2.49​
340​ | 43550.4​ | 43042​ | -1.17​

## Slide 9: Observations

## Slide 10: Uncertainty Analysis

## Slide 11: Error Correlation Structures
 | Langley points | Nights | Spectral
Calibration | C | C | PC
Instrument Noise | I | I | I
Temperature | C | I | C
Aerosol variability | PC | I | C
Errors can be correlated between:
  individual points on the Langley (within one night)
  Individual nights
  Spectral bands
C = correlated; I = independent; PC = Partially correlated

## Slide 12: Model Regression Overview
Based upon the work done by Kieffer and Stone [Kieffer and Stone,2005]​
Minor adaptations to the model formulation discussed with T. Stone
Model is based on the lunar irradiance measurements done at Izaña​
  Historical instrument – 933 : 3 years of irradiance data (+/- 400 acquisitions)​
  New ESA instrument – 1088 : <1 year & 3 months of irradiance data (+/-150 acquisitions) ​

## Slide 13: Lunar Irradiance Model
= model spectral band
= lunar reflectance
= absolute phase angle [radians]
= selenographic latitude of observer [degrees]
= selenographic longitude of observer [degrees]
= selenographic longitude of the Sun [radians]

## Slide 14: Coefficient Regression

## Slide 15: Uncertainty Analysis
Perform Monte Carlo analysis to define total uncertainty ​
Including the uncertainty related to the model regressions​
1000 models by modifying the input​
Introduced perturbation based upon the instrument and measurement uncertainty​
Notes: 𝐸_(𝑖,𝜆)^"True" is the nominal “true” value for the TOA irradiance in spectral band 𝜆 for the 𝑖th observation
𝑅_(𝑖,𝜆) is the error in the observation in spectral band 𝜆 for the 𝑖th observation due to random effects, expressed in relative terms, e.g. as a percentage of the true value
𝑆_𝜆 is the error in the observation that is common for all measurements in this band, expressed in relative terms, e.g. as a percentage of the true value
𝐶 is the error in the observation that is common for all measurements in all bands, expressed in relative terms, e.g. as a percentage of the true value
The error values are unknown; but are draws from a probability distribution with a standard deviation given by the relative uncertainty associated with this effect and with an expectation value (central value) of zero.
𝑅_(𝑖,𝜆) takes a different value for every observation. This comes from random processes relating to the measurement of the TOA irradiance for a particular night. These include instrument noise, instrument temperature changes and atmospheric changes – and relates to the relative uncertainty in the Langley Plot intercept that is estimated above [section 0].

## Slide 16
440 nm
500 nm
1020 nm
1640 nm

## Slide 17: Spectral Model using Apollo spectra

## Slide 18: Comparison Procedure
Model Input
Timestamp of the acquisition​
Position of the instrument/platform (J2000 coordinates)​
Irradiance ​
Using these parameters, the model calculates the geometry parameters required for the comparison:​
Phase angle​
Solar selenographic longitude​
Observer selenographic latitude and longitude​
Distances between Sun, Moon and observer.​

## Slide 19: Proba-V

## Slide 20
BLUE
Proba-V vs LIME
LIME
VITO ROLO
NIR

## Slide 21: Pleiades

## Slide 22: Pleiades vs LIME
LIME
CNES ROLO
BLUE
NIR

## Slide 23: GIRO
Simulation for one year of data ​
Configured with PV spectral response​
Quite large variations observed, most probably due to unrealistic simulated positions​
461 nm
650 nm

## Slide 24: Degree of Linear Polarisation
Skuratov, 2011
500 nm

## Slide 25: Model Development
Obtaining additional measurements (6 year of observation needed to sample the range of solar selenographic latitudinal and longitudinal angles)
Pandora observations to be used to develop spectral interpolation
Provision of an uncertainty associated with each Langley plot and a weighted fit that takes into account those uncertainties
Uncertainties associated with spectral interpolation

## Slide 26: Thank you!
sarah.taylor@npl.co.uk
http://calvalportal.ceos.org/lime