## Slide 1
Update on the Air-LUSI project (to achieve absolute calibration with lunar models)
K. Turpie, UMBC / JCET
GSICS Lunar Calibration Workshop
(on the Ætheric Plane)
16 November 2020

## Slide 2
Air-LUSI Team
Kevin Turpie, PI (UMBC/GSFC 616.2)
Steve Brown, Co-I (NIST)
John Woodward, Co-I (NIST)
Andrew Gadsden, Co-I (U of Guelph)
Tom Stone, Co-I (USGS)
Andrew Newton (U of Guelph)
Steve Grantham (NIST)
Tom Larason (NIST)
Stephen Maxwell (NIST)
Joe Rice (NIST)

## Slide 3
Current ground-based systems are subject to atmospheric challenges, while future on-orbit measurements are subject to on-orbit risks. Air-LUSI addresses these issues with:
High accuracy: <1% (k=1), 415 -1000 nm (target <0.5%)
High spectral resolution: 3.7 nm with 0.8 nm sampling (300 – 1100 nm)
Small atmospheric correction: ~3% from molecular scatter and O3
Calibration checked before and after use.
Calibration monitored up to the point of data collection.
Ready for operation (TRL 8).
Air-LUSI is an essential validation tool, providing confidence to new datasets that nothing else can, and a potentially valuable contribution to a multi-dataset approach to building a new lunar reference.

## Slide 4
ARTEMIS – Autonomous, Robotic TElescope Mount Instrument Subsystem
    Uses tracking camera on telescope and computer controlled PID loop.
    Keeps telescope fixed on the Moon to within 0.1°.
IRIS – IRradiance Instrument Subsystem
    A non-imaging telescope (integrating sphere at focal point).
    Light fed via a fiber optic cable to a spectrograph.
    On-board LED validation source.
    Instrument enclosure keeps the spectrograph and validation source at surface-level P & T during flight.
ARTEMIS Control Assembly
ARTEMIS Control Assembly
Installation
ARTEMIS Mount
IRIS Telescope
ARTEMIS
Tracking Camera
IRIS
Instrument Enclosure
IRIS Fiber Optics
Cable Harness
IRIS Instrument Enclosure
Installation
IRIS Telescope in view port
Pilot Switch Simulator
Air-LUSI Subsystems
Notes: What is your critical path? A bulleted list of steps to our objective. Get h/w ready for engineering test for lunar obs late in May.

## Slide 5
NASA ER-2
Altitude
69,000 ft
21 km
Aft-body
Zenith
Viewport
Superpod
Mid-body
Aft-body
Fore-body
The telescope and robotic mount are in Superpod Aft-body.
The sensor enclosure and control computer are in Mid-body.
IRIS telescope views to port through Aft-body Zenith Viewport.
ABOVE
RIGHT
BELOW
RIGHT
5

## Slide 7
13 Nov 2019 (UTC)
Flight plan for first flight of the Demonstration Flight Campaign

## Slide 8
AFRC Lab
ER-2 (Hangar)
ER-2 (Airborne)
Air-LUSI Lab
Measurement Flow Chart
Pre-Campaign Characterization
Lunar Measurements
Pre-Upload
Calibration
Post-Offload
Calibration
Post-Campaign Characterization
Functionality and System Tests
Descent
Calibration Monitoring
Ascent
Calibration Monitoring
System-Level Tests
System-Level Tests
Pre-Flight
Calibration
Pre-Flight Checks
Last Flight?
No
Yes
Deploy
Return
Upload
Offload
Takeoff
Dark
LED Source
LED Source
Dark
Moon
Functionality Tests
Functionality Tests
Landing

## Slide 9
air-LUSI Flight Campaigns
Engineering Flight Campaign
air-LUSI executed two ~2-hour flights from 1-2 (UTC) August 2018.
Subsystems were successfully tested during several flights.
Moon was observed for 30-40 minutes at ~21 km altitude for two night flights.
Data was recorded the 2nd night at a lunar phase angle of +53°.
Demonstration Flight Campaign
air-LUSI executed five ~2-hour flights from 13-17 (UTC) November 2019.
Moon was observed for 30-40 minutes at ~21 km altitude for each flight.
Flights observed phase angles of +10°, +21°, +34°, +46° and +59°.
Weather was fair for takeoff and landing. Last flight was shifted earlier to accommodate other aviation activities in the area.
Raw data view via telemetry during engineering test flights
13 Nov 2019 (UTC)
Flight plan for first flight of the Demonstration Flight Campaign

## Slide 10
Air-LUSI Initial Results

## Slide 11
ß
Chart: Wavelength [nm]
Initial Error Budget
Air-LUSI Lunar Irradiance Uncertainty Budget

## Slide 12
MODTRAN setup:
  slant path to space from 21.215 km altitude, zenith angle = 22.047°
  1976 US Standard Atmosphere, 380 ppmv CO2
MODTRAN atmospheric transmission for air-LUSI flight 1
O2
Initial Error Budget

## Slide 13
At-sensor uncertainties are <1% (k=1) from 415 nm to 1000 nm.
Over most of the spectral range the uncertainty is dominated by IS Temporal Stability component and we are working to reduce the IS effect.
Next biggest contributor in the uncertainty is the FEL lamp irradiance calibration (which might be cut in half by using SIRCUS before future campaigns).
At the ends of the spectral range, wavelength uncertainty becomes a significant contributor.
Low signal levels in the UV are also a significant contributor.
Atmospheric correction uncertainty contributes a very small amount to the at-sensor uncertainty.
Uncertainty Budget Takeaway
Initial Error Budget

## Slide 14
Air-LUSI is very useful for development of an accurate absolute lunar calibration.
  Not affected by issues of orbiting systems or atmospheric challenges.
  It will be critical to lay the groundwork for current and future efforts.
  We know that it works now.
Air-LUSI data can be used to
  validate other lunar irradiance characterization efforts.
  used directly to improve the absolute accuracy of ROLO.
Demonstration Flight Campaign successfully yielded
  Lunar Spectral Irradiance for five nights at phases: 10°, 21°, 34°, 46° and 59°.
  Error budget give an uncertainty of <0.8 % (k=1), which can be improved.
  We are currently verifying that this uncertainty estimate is accurate.
Summary

## Slide 15
Summary
Air-LUSI is wrapping up data reprocessing and will begin
  Comparing air-LUSI to LIME, ESA data (Tenerife), PLEIADES and ROLO.
  Doing initial work on repairs and maintenance.
  Several tasks dependent on lab access have delayed efforts.
Future flights depend on funding sources, aircraft availability and lunar observational windows.
We are currently considering a campaign as early as this January, however, because of the pandemic we may be delayed until late 2021 or early 2022.

## Slide 16
Thank You!
air-LUSI Team (Left to Right) – Steven Grantham, Andrew Newton, Kevin Turpie, John Woodward, Tom Larason, Stephen Maxwell (not shown: Steve Brown, Andrew Gadsden, Andrew Cataford, and Tom Stone)

## Slide 17
BACKUP SLIDES
17

## Slide 18
To make very accurate, SI-traceable lunar spectral irradiance measurements.
The objective for the Engineering Flight Campaign was <2% (k=1). Current error budget indicates an uncertainty <0.8% (450 – 900 nm) for the Demonstration Flight Campaign.
Air-LUSI expects improved measurement accuracy with each campaign.
Ultimately reach further below 1%, which we believe can be done in 2021.
Air-LUSI Primary Objective

## Slide 19
The ability to use the Moon an extremely stable, highly accurate absolute reference that is available to all Earth observing satellites could better support:
Cheaper Satellites
Faster Satellite Calibration
Intercalibration of Earth Observing Systems (US govt, International and Commercial)
  Interoperability
  Time Series Accuracy
New Scientific Information
  Better Understanding of Atmospheric Correction over Water
  Better Calibrated NIR Bands
Why absolute lunar calibration?

## Slide 22
Robotics
Air-LUSI network
and telemetry
Instrument
Function
Mobile Pilot
Communication
air-LUSI “Mission Control”

## Slide 23
ISS
Transfer Standard Spectrograph (TSS)
IRIS Spectrograph
IRIS Telescope
DTSS
DIRIS
ETSS
EIRIS
Air-LUSI is calibrated by observing a lamp-illuminated integrating sphere source (ISS) with the approximate angular subtense of the full moon. The output of the ISS is measured by a transfer standard spectrograph (TSS) that holds an SI-traceable scale form NIST.
The inverse square law and measurements of the distance between the integrating sphere and the transfer standard and IRIS are used to transfer the scale to IRIS.
ISS
Mount
Alignment laser
E-head mount
Linear stage
IRIS telescope
ISS
Calibration Scheme

## Slide 24
FEL Lamp Irradiance
Transfer Standard Spectrograph
Integrating Sphere Source Irradiance
air-LUSI Irradiance Responsivity
Lunar Irradiance @ air-LUSI
High Temperature Blackbody
MODTRAN
Top-of-the- Atmosphere Lunar Irradiance
ROLO
5
4
3
2
1
air-LUSI Calibration/Lunar Measurement Chain
6
E – Irradiance, S – Signal, D – Distance, λ – wavelength.
Subscripts are the measuring instrument: Transfer Standard Spectrograph (TSS) or IRIS.
Superscripts are light source: FEL, Integrating Sphere Source (ISS), or Moon.
Measurement Equation (w/o Atmospheric Correction)

## Slide 25
The exo-atmospheric lunar spectral irradiance ELUSI is computed by dividing the at-sensor lunar spectral irradiance by the transmission Ts as a function of λ. The relative uncertainty is RSS with the at-sensor uncertainty of the transmittance.
Error propagation for atmospheric correction:
≈ 0.20% @λ = 400 nm
≈ 0.32% @λ = 620 nm
Assuming a conservative 10% error in the model, maxima are given at the following wavelengths, which has a nearly insignificant impact on the at-sensor uncertainty.
Initial Error Budget

## Slide 26
Address Satellite / ROLO Discrepancies – Air-LUSI observations should help determine why calibrated satellite measurements and ROLO differ by 5-12 %.
Lunar Reference Model Improvement – Air-LUSI can provide SI–traceable calibration data that could be used to improve or validate current lunar reference models (e.g., ROLO, GIRO, RIMO and LIME).
Lunar Calibration Methodology Improvement – Air-LUSI can discern what portion of these differences come from ROLO model bias and satellite lunar measurement methodologies.
Air-LUSI Data Initial Application
Calibrated Satellite Measurements
ROLO Predictions
Air-LUSI Measurements
How close to truth is it?
Adjust
ROLO
Inform
Satellite Cal Community
Truth (± unc)
ROLO differs from Truth
Sats differs from Truth
Notes: Understanding the ROLO Uncertainty

## Slide 27
Other data sources can validate against Air-LUSI data.
air-LUSI Measurements
ESA Tenerife Measurements and LIME
ARCSTONE
NIST MLO-LUSI
Air-LUSI Data Future Application
1.
2.
4.
3.
Air-LUSI is ready.
ESA – currently 150+ nights of multi-band data.
NIST MLO-LUSI – ready in several months.
ARCSTONE – ready in a few to several years.
Possible future comparisons
Aqua MODIS
NPP and JPSS VIIRS
Landsat 8 & 9 OLI
GEOS ABI
PACE
SBG
Other possible near-term comparisons:
CNES PLEIADES
Mazarico Photometric Model
Terra MODIS

## Slide 28
air-LUSI data and other data sources can then be used to improve Lunar Reference model incrementally.
air-LUSI Measurements
ESA Tenerife Measurements
ARCSTONE
NIST MLO-LUSI
COMMON
LUNAR IRRADIANCE MODEL
1.
2.
3.
air-LUSI is ready.
ESA – currently 150+ nights of multi-band data.
NIST MLO-LUSI – ready in several months.
ARCSTONE – ready in a few to several years.
4.
air-LUSI Data Future Application
air-LUSI data can be used to initially improve the model in order to inter compare datasets.