## Slide 1
MSG/SEVIRI radiometric performance monitoring and end-of-life tests using lunar calibration
S. Wagner, V. Debaecker*, OPS team
* Telespazio support

## Slide 2: MSG SEVIRI
SEVIRI warm channels
Spinning satellite
15min repeat cycle in full disc configuration
5min repeat cycle in rapid scan mode over Europe

## Slide 3: SEVIRI warm channel calibration
SEVIRI does not feature a VIS-NIR on-board calibration device and relies on desert vicarious calibration by RSP/INRC for updating the Level 1.5 calibration coefficients
SEVIRI Solar Channel Calibration system (SSCC), developed for MSG1 launch in 2002 by Y. Govaerts & M. Clerici (TGRS 2004), has been used to extract calibration coefficients
SSCC is based on vicarious bright desert calibration (over 18 targets) plus dark ocean targets (9 targets) for verification purposes
The RTM simulations of Top-Of-Atmosphere radiances are compared with TOA measured signal and absolute calibration coefficients and temporal drift are estimated
Location of targets used in SSCC
SSCC calibration example

## Slide 4: MSG/SEVIRI lunar intrusions
Low-res observation and Moon detection
High-res observation and Moon detection

## Slide 5: MICMICS Workflows
Instruments
Instruments
Lunar Calibration
Desert Calibration
DCC Calibration
Lunar Inter-Calibration
Inter-Calibration Blender
DCC Inter- Calibration
Rayleigh Calibration System
Rayleigh Inter-Calibration System
Imager-Hyperspectral Inter-Calibration
NWP Calibration
Imager-Imager Inter-Calibration
ICBS
IHICS
IIICS
NWPC
Ray-Matching Inter-Calibration System
NWP Inter-Calibration
NWPIC
DCC Extraction Module
Lunar Extraction Module
DECS
DCCICS
DCCCS
LCS
LICS
RAYMICS
RAYCS
RAYICS
DCCXTRM
LXTRM
MICMICS
CPMs
VIS/NIR channels
MSG SEVIRI VIS/NIR
S3
SLSTR & OLCI VIS/NIR
MTG LI VIS
EPS AVHRR VIS/NIR
EPS-SG 3MI VIS/NIR
MTG FCI VIS/NIR
Thermal channels
MSG SEVIRI IR
MTG FCI IR
S3 SLSTR IR
EPS AVHRR IR
EPS-SG VII IR
EPS-SG MWI/ICI MW
CO2M
CLIM & MAP
EPS-SG VII VIS/NIR
Pre-Colocation Tool
Colocation Tool
Subsetter Tool
Data-Fusion Tool
SUBTL
DFTL
PCOLT
COLT

## Slide 6: Preparation of the input to the GIRO / LESSSR
  Oversampling = 1
      Will be revisited in the context of GEO imagery
  Solid angle supported by the detectors:
      Derived from the overall FOR extent (18.4 deg x 18.4 deg) and the number of detector acquisitions
  Observed lunar irradiance
  Current approach:
    Extremely good noise levels for SEVIRI  Integration of the whole imagette

## Slide 7: Some examples of time series…
Impact on the drift estimate of a few occasional observations only…
SEVIRI operational calibration (in the) = step function (corresponding to the calibration updates)
For SEVIRI monitoring , calibration of the instrument counts on the GIRO scale
 Isolate the instrument performance from the on-ground calibration process (method + updates)

## Slide 8: MSG/SEVIRI monitoring
Official results provided by MPSTAR  table from last EUMETSAT State of Observing System Report for 2022

## Slide 9: Implementation of the LESSSR model
LESSSR = based on SCIAMACHY lunar measurement + RELAB data
SCIAMACHY measurements between -80 and 20 deg lunar phase
LESSSR model  disk-equivalent lunar reflectance at 1AU at a grid of 5nm between 300nm and 2500nm:
Pa = P0 + P1 g’ + P2 g’ 2 + P3 g’ 3 with and g0 = -30 deg (SCIA ref. angle)
Pb = P4 φ + P5 θ + P6 Φ
Pc = P7 * exp(-|g| * P10) + P8 * exp(-|g| * P11) + P9 * cos((|g|-P12) * P13)
Rlunar = exp(Pa + Pb + Pc)
With g the lunar phase, φ and θ the selenographic longitude and latitude of the observer, and Φ the selenographic longitude of the Sun.

## Slide 10: MSG1/SEVIRI – Irradiance biases

## Slide 11: MSG1/SEVIRI – Irradiance biases

## Slide 12: MSG1/SEVIRI – Irradiance biases

## Slide 13: MSG1/SEVIRI – Irradiance biases
Work still needed on the current implementation of the LESSSR applicator
Additional model to complement GIRO in MICMICS
SCIAMACHY provided hyperspectral measurements from space

## Slide 14: Meteosat-8 end of life tests
Goal of the EOL test
  Attempt to verify the medium term (=24h) radiometric stability requirement on SEVIRI warm channels (= 0.1% of dynamic range/day, outside eclipse days and vicarious calibration and excluding variation due to changing illumination angle throughout the day)
  Never verified as vicarious calibration cannot achieve the accuracy / uncertainties needed over 24h
  Check the possibility to perform de-saturated lunar acquisitions with TIR channels.
Formally, verification of the requirement is currently not possible due to the current uncertainties of the GIRO model (0.5% – 1% over long time series or with identical lunar phases).
But lunar models will be improved, when measurements are only possible in specific times…

## Slide 15: Meteosat-8 end of life tests
During MSG4 commissioning
    Depointing (North / South)
  Limitation = uncertainties on the GIRO for relative calibration (0.5% - 1%)
Two test sequences:
  Each test = 8 hours long
  One image every 5min
  Test 1: 09.09.2022, normal gain / dedicated rapid scan in the Southern hemisphere
  Test 2: 12.09.2022, modified gain to try to de-saturate (some) IR channels (12.0 and 13.4m) / nominal RSS over the Northern hemisphere

## Slide 16: Meteosat-8 end of life tests

## Slide 17: Meteosat-8 end of life tests

## Slide 18: Meteosat-8 end of life tests

## Slide 19: Meteosat-8 end of life tests

## Slide 20: Meteosat-8 end of life tests

## Slide 21: HRV sequence – 09.09.2022 (including retrace)

## Slide 22: HRV sequence – 09.09.2022 (sliced)

## Slide 23: Results – relative calibration on the GIRO scale
-16.7 < Phase angle < -8.4
Variability over ~8h goes beyond the stated requirement
Not fully understood yet what is the root cause, but possibly some explanations
Inherent model uncertainties shall not be forgotten (~0.5% - 1% over long time series)!

## Slide 24: Results - analysis
Over the phase range covered by the test  variability ~1.5% due to the dependency in the NIR1.6
VIS0.6 + VIS0.8  variability observed over the MSG1 lifetime do not allow the identification of a dependency
But a smaller effect observed with AHI, at least for VIS0.6
red = linear fit
blue = polynomial fit (n=4)
green = exponential function fit
Uncorrected
Corrected
Polynomial deg. 4

## Slide 25
Band1 (0.47μm) Band2 (0.51μm) Band3 (0.64μm)
Band4 (0.86μm) Band5 (1.6μm) Band6 (2.3μm)
Himawari-8/AHI lunar phase angle dependence in ΔIrr
ΔIrr (%)
Difference of lunar irradiance between the observation and GIRO [%]
ΔIrr (%)
Lunar phase angle [deg]
Lunar phase angle [deg]
Lunar phase angle [deg]
MSG1/SEVIRI NIR1.6
MSG1/SEVIRI VIS 0.6
SEVIRI results from B. Viticchie at the Lunar Calibration Workshop in Dec. 2014
Good agreement with SEVIRI NIR1.6, but not with SEVIRI VIS 0.6
Moon pixel selection w/ larger DC threshold (~30)
Presented by M. Takahashi (JMA)

## Slide 26
Band1 (0.47μm) Band2 (0.51μm) Band3 (0.64μm)
Band4 (0.86μm) Band5 (1.6μm) Band6 (2.3μm)
Himawari-8/AHI lunar phase angle dependence in ΔIrr
ΔIrr (%)
Difference of lunar irradiance between the observation and GIRO [%]
ΔIrr (%)
Lunar phase angle [deg]
Lunar phase angle [deg]
Lunar phase angle [deg]
MSG1/SEVIRI NIR1.6
MSG1/SEVIRI VIS 0.6
SEVIRI results from B. Viticchie at the Lunar Calibration Workshop in Dec. 2014
Good agreement with SEVIRI NIR1.6, but not with SEVIRI VIS 0.6
Moon pixel selection w/ lower DC threshold (=25) + Connected-component labeling
Presented by M. Takahashi (JMA)

## Slide 27: Results - analysis
Over the phase range covered by the test  variability ~1.5% due to the dependency in the NIR1.6
VIS0.6 + VIS0.8  variability observed over the MSG1 lifetime do not allow the identification of a dependency
But a smaller effect observed with AHI, at least for VIS0.6
Others aspects than are under investigations:
  Offset correction
  Effect of the distance to the Moon and potentially some oversampling
red = linear fit
blue = polynomial fit (n=4)
green = exponential function fit

## Slide 28: Results - analysis
Over the phase range covered by the test  variability ~1.5% due to the dependency in the NIR1.6
VIS0.6 + VIS0.8  variability observed over the MSG1 lifetime do not allow the identification of a dependency
But a smaller effect observed with AHI, at least for VIS0.6
Effect of the offset correction on the data
  SEVIRI noise levels = very low
  Currently integration done on the complete imagette
Effect of the distance to the Moon and potentially some oversampling  currently under discussion with GEO/GF

## Slide 29: Results - analysis
Over the phase range covered by the test  variability ~1.5% due to the dependency in the NIR1.6
VIS0.6 + VIS0.8  variability observed over the MSG1 lifetime do not allow the identification of a dependency
But a smaller effect observed with AHI, at least for VIS0.6
Effect of the offset correction on the data
  SEVIRI noise levels = very low
  Currently integration done on the complete imagette
Effect of the distance to the Moon and potentially some oversampling  currently under discussion with GEO/GF
Jumps observed when Tback Off is changed  change in the scanned sector  Caused not known yet

## Slide 30: Meteosat-8 end of life tests
Test 1 conducted successfully despite an initial anomaly (related to the set-up of the measurement triggering ) which reduce slightly the number of observations
Current analysis of the data shows potential for improving the end results
Test 2 sequence to be further investigated after some issues with the measurements
The anomalies we faced during the two tests will allow a better definition of similar tests in the future.

## Slide 31: VIS0.6 sequence – 12.09.2022 (sliced)
Offset reset due to the presence of the Earth in the offset angular window…
 Manual reprocessing need to adjust the count levels

## Slide 32: IR13.4 sequence – 12.09.2022 (sliced)
Wrong initial IR gain set-up  corrected later on during the acquisitions
Two sets of gains implemented  only one was ok!
IR12.0 + IR13.4 bands unsaturated for part of the observations
  Further data analysis needed
  Possibility to collaborate with Martin Burgdorf (university of Hamburg)

## Slide 33: Future developments
Revisit the solid angle + oversampling calculation for SEVIRI and geostationary missions in general
Finalise the implementation of the LESSSR model in MICMICS
  To complement systematically the GIRO simulations, using same input / output file format
  Analyse further the data acquired during the MSG1 end-of-life tests
   Check the possibility to use the IR measurements
   Investigations on the HRV data

## Slide 34
Thank you!
Questions are welcome.

## Slide 35
Back-up slides

## Slide 36: MSG/SEVIRI lunar intrusions
Low-res observation and Moon detection
High-res observation and Moon detection

## Slide 37
MTF post-launch assessment using lunar imagery

## Slide 38: Other usages of lunar imagery – MTF post-launch assessment

## Slide 39: Other usages of lunar imagery – MTF post-launch assessment

## Slide 40: Other usages of lunar imagery – MTF post-launch assessment
Example of MTF determination using the Moon edge
Method similar to approach used by Truman & Xiong on ABI (SPIE 2019, DOI: 10.1117/12.2528138)
Spatial frequency
Normalised MTF
Still work in progress
Impact of the moon surface variability to be accounted for the MTF retrieval

## Slide 41: Other usages of lunar imagery – MTF post-launch assessment
From Truman & Xiong, SPIE 2019 – MTF for GOES-16/ABI

## Slide 42
Using lunar thermal infra-red observations…

## Slide 43: AHI measurements TIR
From: Utilization of a meteorological satellite as a space telescope: the lunar mid-infrared spectrum as seen by Himawari-8
a–e Example images of quality level (QL) 1–5. QL of 1, 2, 3, 4, and 5 correspond to the Moon hidden more than half, cut by swath boundaries, hidden partly by the Earth, partly out of the scanned area, and fully captured, respectively. The images were retrieved from the band-13 data taken at 2015/04/01 00:40, 2016/10/26 12:00, 2016/11/23 09:30, 2016/12/09 23:20, and 2017/10/30 23:10. f The histogram of the number of QLs for each year from 2015 to 2021. The gray, black, green, blue, and red bars correspond to QL of 1, 2, 3, 4, and 5, respectively
Band-13 = IR 10.4

## Slide 44: AHI 3.9 band (B07)
Super Moon
Himawari-8
Example of Moon acquisition by B07