## Slide 1: Straylight Corrections to Improve ABI Lunar Irradiance Measurement: A Case Study of GOES-16/17 ABI Spatial Uniformity Validation
Fangfang Yu1 and Xiangqian “Fred” Wu2
Acknowledgements: GOES-R CWG team for the moon appearance prediction and lunar image processing, GOES-R PRO for the coordination effort, GOES-R MOST and operation teams for the lunar image collections, and EUMETSAT for sharing the GIRO model
4th Joint GSICS/IVOS Lunar Calibration Workshop, 4-8 December 2023
1: University of Maryland
2: NOAA/NESDIS/STAR

## Slide 2: Introduction
The ratio between the measured and predicted lunar irradiance is often used to study the instrument performance
  In-orbit calibration (e.g. solar calibration) validation
  Long-term trending
  Cross-calibration
  Instrument characterization
In addition to the accurate irradiance prediction, accurate satellite measurement is also crucial for uncertainty reduction and accurate interpretation of the results.
This talk is an example to improve the lunar irradiance measurements with straylight corrections to characterize the ABI spatial uniformity in the East-West (EW) direction
Reference:
  Yu, F., X. Wu, X. Shao, and H. Qian: “Characterization of the east-west spatial uniformity for GOES-16/17 ABI bands using the Moon,” Remote Sensing, 2023, 15, 1881, https://doi.org/10.3390/rs15071881.

## Slide 3: GOES-R ABI Instrument
ABI Bands: 16 bands
  6 visible and near-infrared (NVIR) bands
  10 infrared (IR) bands
Two independent scan mirrors
  North-South (NS)
  East-West (EW)
On-orbit calibration for all the bands
  On-orbit solar diffuser (SD) for VNIR bands
  Blackbody for IR bands
GOES-17 ABI, due to its cryocooler misfunction, has different instrument calibration configurations from the original design, which are set for GOES-16/18.
ABI Optical Architecture. The FPM temperatures are set for GOES-16/18. For GOES-17, there is no thermal control for the VNIR FPM temperature and the IR FPM temperatures are elevated to higher values.
Figures are from GOES-R Series Data Book
ABI scanner

## Slide 4: GOES-R ABI Lunar Scans
GOES-16 ABI Mode 3 Timeline
Each normal ABI timeline consists of full-disk scans, CONUS scans, MESO scans and calibration targets’ scans
The lunar images are collected with MESO scans which consists of two vertically aligned swaths
  The Moon is scanned with only one of the two swaths
Lunar image collections for Response Versus Scan-angle (RVS) study
  Chasing the Moon from west to east in the space
    high angular resolution
    Combined data cover the full scan angle range for the FD images
  5o < |Lunar phase angle| < 60o, in general
G16 ABI Channel 1 (0.47µm) image scanned with one MESO swath
Earth
Moon trajectory within ABI field of regard

## Slide 5: ABI VNIR Radiometric Calibration
Time-series of the NS scan angles for the spacelook scenes of this timeline.
Locations of the G17 spacelook events (in red dots) for the timeline between 20:02-20:17Z on 07/30/2018.
GOES-16 ABI Mode 3 Timeline
The lunar chasing images for GOES-16/17 RVS study were all collected with ABI Mode 3 timeline

## Slide 6: Data Collections for the RVS Study
Satellite | GOES-16 |  |  |  | 
Date | Duration (UTC) | Phase Angles | NS Scan Angles | EW Scan Angles | FPM Temp. Change
2017/02/11 | 18:19 - 19:17 | +9.1o - +10.8o | +9.5o - +9.4o | -6.5o – +4.7o | 0.1 K
2017/02/12 | 18:54 - 19:04 | +21.0o - +21.4o | +5.9o - +5.9o | -10.1o - -9.0o | < 0.1 K
 | 20:12 - 20:22 | +23.4o - +23.8o | +5.7 o - +5.7o | +7.3o – +7.6o | 
2017/02/14 | 20:20 - 20:27 | +44.9o - +45.1o | -1.7o - -1.7o | -10.5o - -11.7o | < 0.1K
 | 21:52 - 22:00 | +47.6o - +47.9o | -1.9o - -1.9o | +8.2o - +10.8o | 
2017/04/12 | 18:57 - 19:38 | +17.0o - +18.7o | -9.5o - -9.4o | -3.6o - + 5.0o | <0.1 K
2017/07/12 | 20:32 – 21:16 | +41.5o - +42.8o | -10.2o – 10.1o | -5.1o - + 3.6o | < 0.1 K
Satellite | GOES-17 |  |  |  | 
Date | Duration (UTC) | Phase Angles | NS Scan Angles | EW Scan Angles | FPM Temp. Change
2018/07/30 | 19:50 – 20:45 | +32.3o - +33.9o | -9.6o - -9.4o | -6.5o - + 5.3o | 0.1 K
2018/08/28 | 18:53 - 19:01 | +25.0o - +25.3o | -3.2o – -3.1o | -10.1o - -11.7o | 0.4 K
 | 20:21 - 20:28 | +27.7o - +28.0 o | -2.9o - -2.9o | +7.7o - + 9.3o | 
2018/08/29 | 19:37-19:42 | +36.6 o - +36.9o | +1.0o – 1.0o | -11.8o - -10.5o | 0.5 K
 | 21:07-21:16 | +39.5o - +39.7o | +1.3o -+1.3o | +8.1o - + 8.5o | 
2018/10/20 | 14:27 – 15:15 | -48.1o - -46.7o | -10.0o- -9.8o | -3.3o - + 6.4o | 0.9 K
Each chasing event lasts from a few minutes to about one hour
Lunar phase angle changes in less than 2 degrees
Stable VNIR FPM temperature for G16, yet experienced FPM temperature variation for G17

## Slide 7: Subset ABI Lunar Images
B01
B02
B03
B04
B05
B06
Band | Central wvlen (um) | Columns | Rows
1 | 0.47 | 800 | 636
2 | 0.64 | 1600 | 1380
3 | 0.87 | 800 | 636
4 | 1.38 | 400 | 352
5 | 1.61 | 800 | 636
6 | 2.25 | 400 | 352
Subset image size:
Scheme for the best detector selected

## Slide 8: ABI Lunar Irradiance Calculation
No definitive boundary
  Size of the subset image?
Ghosting at B06
  Consistent for all B06 lunar image
  Negligible for the total lunar irradiance
Noise
  Negligible for irradiance calculation
Ω: Sample solid angle project of IFOV
Oversampl_fa: oversampling factor
Radi,j : calibrated radiance for the sample at (i,j) image coordinate
Subset lunar images with lunar samples (albedo > 0.33%) masked as -0.1%.
Coherent noise, within ABI noise requirement
Ghosting, within ABI radiometric cal requirement
Originates from the Moon, thus should be accounted in the lunar irradiance calculation

## Slide 9: Impacts of Model Uncertainty
RVS is characterized with the irradiance ratio from the chasing events
Impacts of the model uncertainty:
  Relatively stable ratio within each event
  Large gaps among the chasing events for the long wavelength channels (e.g. B04 – B06)
Lunar irradiance ratio versus EW scan angle for G16 ABI

## Slide 10: Normalized Ratio for RVS Characterization
Normalized ratio to overcome the bias among the chasing events

## Slide 11: Initial Results
Relatively large variation within each event at G16 B01-B02 and G17 B01 – B03
  B01: up to 2% variation
  Impact of earthshine straylight on B01-B03
B04-B06: within 0.5%

## Slide 12: Straylight at Spacelook Scenes
Straylight may occur at the spacelook scenes near the polar regions around the satellite noon time
B01 displays the largest magnitude, followed by B02 and B03, becoming in-distinguishable from B04 – B06
  Origin: light scattered from the Earth
At B01, the straylight can elevate the spacelook by 1-2 counts out of the 14-bit depth of the ABI calibration data
  Negligible calibration error at sample level: <0.01% albedo for a typical 40% albedo of the Earth
Time-series of the mean spacelook value for G16 during the lunar chasing event on 2017/02/11. The vertical dash line is the starting time for each 15-minute timeline.

## Slide 13: Spacelook Correction for B01-B03
Stable FPM data
where is the detector mean spacelook value from the spacelook events at the Equator in a timeline.
Unstable FPM data for G17
where is estimated with the linear regression between the FPM temperature and spacelook count at the spacelook observation time
Time-series of the mean spacelook value for G17 VNIR bands during the lunar chasing event on 2018/10/20. The gray line is referred to the second y-axis for the FPM temperature. The vertical dash line is the starting time for each 15-minute timeline.

## Slide 14: Results of Spacelook Correction for B01-B03
Removed the scattered outliers in the long chasing events
Improved the consistency of normalized irradiance ratios
Common variation patterns remained: elevated normalized ratio when it approaching to the illuminated Earth limb
  Straylight spilled into the space from the illuminated Earth

## Slide 15: Magnitude of Spilled Light
Spilled light at ~1 count can be detected at about 0.5 degrees from the Earth limb.
The small elevation of the delta count at ~1 count can cause an extra radiance of ~ 0.058 W/m2Srµm or 0.009% albedo at each sample.
This error can contribute to up to ~1% B01 irradiance variation
  Similar magnitude as impacted with the contaminated spacelook scenes
Albedo value for the G17 FD equator swath (swath #11) collected on 2018/07/26 at 19:15 UTC when the spacecraft was in a rolling position for special scans.
Variation of averaged delta count with the scan angle for the samples in north half-swath shown above

## Slide 16: Spilled Light Correction
where is the radiance averaged from the space samples at two edges of a lunar image, and is the spacelook-corrected lunar sample radiance of the sample in the lunar image

## Slide 17: Final Result after Straylight Corrections
RVS in the EW direction for G16 (left) and G17 (right)
RVS is within 0.5% variation for all the G16/17 VNIR bands
  Straylight corrections were applied to B01-B03

## Slide 18: Summary
Accurate interpretation of the instrument performance with lunar irradiance calibration requires accurate measurement of the lunar irradiance
This could be specifically important for the high lunar resolution images
  Negligible radiance error can be accumulated to a significant level for irradiance value
ABI is a very sensitive instrument
  Capable of distinctive detection of a small response change at ~1 count out of the 14-bit depth data.
  Two types of weak straylight originated from illuminated Earth are detected and corrected in this study to improve the irradiance measurements
    Light scattered on the spacelook scenes
    Light spilled into the vicinity space area where the Moon may appear
After the straylight corrections, the RVS variations for all the G16/17 VNIR bands are consistent and within 0.5%.

## Slide 19: Why Significant Irradiance Variation?
Irradiance measurement:
Number of lunar image samples: ~1.5x105 to 2.2x106 per channel
Lunar irradiance can be sensitive to small calibration error which has negligible impact on sample radiance
Radi,j : calibrated radiance at (i,j) image coordinate in a subset lunar image.
Impacts of uncertainty of solid-angle (Ω) and oversampling factor are cancelled out in the normalized irradiance ratio in this study.
Band | Columns | Rows
1 | 800 | 636
2 | 1600 | 1380
3 | 800 | 636
4 | 400 | 352
5 | 800 | 636
6 | 400 | 352
#samples = 1.48E+5 – 2.2E+6

## Slide 20: Discussion: Straylight in Lunar Irradiance Calculation
Straylight is the light that reaches the detector but originates outside the field of view
Depending on the sources, the straylight should be accounted or extracted from the lunar irradiance measurement. For ABI,
  If originated from the Moon =>
    need to be accounted in the lunar irradiance measurement
  If originated from the illuminated Earth or Earthshine =>
    Need to be corrected for the accurate lunar irradiance measurement

## Slide 21: Recommendations for Lunar Image Collections
Schedule: use the MESO scans near the middle time in a Mode 6 timeline, if possible
  Spacelook near the Equator, thus not affected with the possible scattered light near the polar region
Collection: the Moon should be collected at least certain distance from the Earth limb or at the dark side to the Earth
Subset lunar images: cover the moon and the vicinity area where lights is scattered or leaking from the Moon, but ensure the area is not too large so as to amplify the potential calibration error in the irradiance values.

## Slide 22: Declaimer
The scientific results and conclusions, as well as any views or opinions expressed herein, are those of the authors and do not necessarily reflect those of NOAA or the Department of Commerce.