## Slide 1
Truman Wilsona and Xiaoxiong Xiongb
aScience Systems and Applications, Inc.
bSciences and Exploration Directorate, NASA/GSFC
Fourth Joint GSICS/IVOS Lunar Calibration Workshop, 4-8 December 2023, Darmstadt, Germany
Surface Corrected Lunar MTF Measurements in MODIS and VIIRS

## Slide 2
Outline
2
Background
Spatially Resolved Lunar Model
  SP Model provided by Toru Kouyama (AIST/JAPAN) in Jan. 2021.
Projection of the Model Images
Correction of the Edge-Spread Function (ESF)
MTF Trending for MODIS/VIIRS

## Slide 3
MTF From the Moon
3
Lunar limb simulates a laboratory knife-edge measurement.
Moon motion and curvature allows for sub-pixel sampling.
Issue: Lunar surface variation complicates construction of the edge-spread function (ESF).
Solution: Use a spatially resolved lunar model (SP Model from AIST* in Japan) to remove surface variation.
Most work looks at higher resolution Moon images (e.g. GOES ABI). Can it work at low resolution?
Lunar Surface
Deep
Space
Edge
*Provided by Toru Kouyama, Jan. 2021
Notes: MODIS combines EV imagery with OBC observations each scan.
Main RSB calibrator is the SD.
Degradation of SD measured by the SDSM.
SV primarily used for BG, but can view the Moon
<click>
The Moon is at a different AOI, used for RVS.

## Slide 4
Instrument Background
4
MODIS
On Aqua and Terra
  Aqua: Afternoon orbit, Terra: Morning orbit
250-m, 500-m, and 1-km resolution bands
Views the Moon through the space-view port (SV)
  Historic phase of -55.5° for Aqua, +55.5° for Terra
VIIRS
On SNPP, NOAA-20, and NOAA-21
  All in the same orbit
375-m and 750-m
  Some unaggregated Moon views with 3x resolution in the scan direction.
Views the Moon through the space-view port (SV)
  -50.5° target phase
Notes: MODIS combines EV imagery with OBC observations each scan.
Main RSB calibrator is the SD.
Degradation of SD measured by the SDSM.
SV primarily used for BG, but can view the Moon
<click>
The Moon is at a different AOI, used for RVS.

## Slide 5
Lunar Image Comparison
5
01
250-m
Terra
Aqua
03
500-m
19
1-km
SNPP
N20 (N21 similar)
I1
375-m
M1
750-m
M6
750-m
VIIRS
MODIS
Notes: MODIS combines EV imagery with OBC observations each scan.
Main RSB calibrator is the SD.
Degradation of SD measured by the SDSM.
SV primarily used for BG, but can view the Moon
<click>
The Moon is at a different AOI, used for RVS.

## Slide 6
SP Model
6
SP Model provided by Toru Kouyama of AIST (Japan) in January 2021.
Derives spatially-resolved lunar radiance predictions across 160 channels from 0.52 to 1.64 μm.
  6 nm resolution from 0.52 and 0.94 μm.
  8 nm resolution from 0.94 and 1.64 μm.
0.5° angular resolution in Selenographic coordinates. (720 x 360 x 160 GeoTIFF output)
Band specific maps are created using a weighted average of the GeoTIFF layers over the RSR.

## Slide 7
Lunar Geometry Definitions
7
The Selene Profiler (SP) model is used to account for geometric differences in the lunar observations.
  The geometric differences can be described using 6 parameters.
  Selenographic Coordinates refer to locations (projections) on the lunar surface.
Parameter | Definition
dSat-Moon | Distance between the satellite and the Moon
dSun-Moon | Distance between the Sun and the Moon
Libration Angles | 
 | Selenographic Longitude of the Observer
 | Selenographic Latitude of the Observer
 | Selenographic Longitude of the Sun
 | Selenographic Latitude of the Sun
Parameter | Definition
dSat-Moon | Distance between the satellite and the Moon
dSun-Moon | Distance between the Sun and the Moon
Libration Angles | 
 | Selenographic Longitude of the Observer
 | Selenographic Latitude of the Observer
 | Selenographic Longitude of the Sun
 | Selenographic Latitude of the Sun
Example shadow image produced with the SP Model developed by AIST, Japan.
The angle between the observer and solar coordinates is the lunar phase angle. (Aqua MODIS observation at -50.3°, right)

## Slide 8
SP Model Projection
8
N20 VIIRS
Aqua MODIS
The corner of each “view” pixel in the imagery is mapped to the SP model output.
For a pixel corner location ():
z is track, y is scan, and x normal to the image plane. and are the measured Moon center.
Invalid values of x indicate deep space.
  A pixel with mixed valid/invalid pixels is an edge.

## Slide 9
Making a Model Image
9
Each pixel corner is projected onto the model map.
Edge pixels contain invalid points, and those corners are not drawn below.
Surface pixels use the mean of model values contained in the pixel area.
Measured Image
Pixel Type
Model Projection
Deep Space
Moon Edge
Moon Surface

## Slide 10
Making a Model Image
10
Each pixel corner is projected onto the model map.
Edge pixels contain invalid points, and those corners are not drawn below.
Surface pixels use the mean of model values contained in the pixel area.
Measured Image
Model Projection
Measured/Modeled Ratio

## Slide 11
Line Profile Correction Example
11
The measured image profiles quickly go from deep-space to a high degree of surface variation across the lunar limb.
Measured Image
Modeled Image
The model data is scaled to the measured data using the ratio of the sums of the measured and modeled data (excluding 0’s).
The measured data is divided by the model wherever valid values of the model exist.
Edge Pixel
Surface Pixel

## Slide 12
Profile Alignment
12
Derivatives of each line profile are used for alignment at the peak.
The model corrected data puts each profile on the same scale at the surface.
The corrected data is binned to the 0.25-pixel level.
  Outside of ± 3.5 pixels, the binned line profile is padded with 0’s and 1’s for the MTF calculation.
Terra MODIS
Band 1

## Slide 13
MTF Calculation Example: MODIS
13
Terra - 01
Aqua - 01
The binned profiles represent the edge-spread function (ESF).
The derivative of the ESF gives the line-spread function (LSF).
The MTF is derived using a fast-Fourier Transform (FFT) of the LSF.
ESF
LSF
MTF
ESF
LSF
MTF

## Slide 14
MTF Calculation Example: VIIRS
14
NPP – M10
N20 – M10
The binned profiles represent the edge-spread function (ESF).
The derivative of the ESF gives the line-spread function (LSF).
The MTF is derived using a fast-Fourier Transform (FFT) of the LSF.
For N20 (and N21), the lunar views are in 1:1 aggregation, so the scan direction is at higher resolution and has a lower MTF value.
ESF
LSF
MTF
ESF
LSF
MTF

## Slide 15
Edge Curvature Correction
15
More profiles fill out the ESF, but come at a decreased edge sharpness due to edge curvature.
Edge curvature correction applied by multiplying the pixel position by .
Curvature correction reduces event-to-event noise.
𝜃
Aqua Band 1 - Track

## Slide 16
Time Series Analysis - MODIS
16
Terra
Aqua
Band 1
Band 1
Band 4
Band 4
Band 11
Band 11
MODIS includes the spectroradiometric calibration assembly (SRCA). Provides scan direction MTF measurements.
Bands 1 and 2 in MODIS show a downward trend in the MTF in the scan direction.
The 1-km bands show increased noise, probably due to the difficulty in matching the model profiles.

## Slide 17
MTF Comparison - MODIS
17
The track direction MTFs are generally higher than the scan direction.
The lunar derived MTF is in reasonable agreement with the pre-launch (PL) data except for the Terra 1-km bands.
The average for each band is still above spec (0.3 at fN). However the 250-m bands have a downward trend in the scan direction.
Terra
Aqua

## Slide 18
Time Series Analysis - VIIRS
18
N20
N21
NPP
Band I1
Band I1
Band I1
Band M7
Band M7
Band M7
Band M10
Band M10
Band M10

## Slide 19
Conclusions
19
The SP model provides a good correction to the lunar surface variation for the higher resolution bands.
  For MODIS, the 1-km bands show higher noise.
  Future VIIRS instruments should have higher resolution like N20/N21 if the lunar observations are taken in the 1:1 aggregation mode.
Surface and edge curvature correction allows for better sub-pixel sampling of the ESF.
The results for most bands are in good agreement with pre-launch assessments and the results from other Moon-based studies.
  For the scan direction, the better sampled ESFs reduce variation in the MTF significantly. However, the 1-km band results for MODIS show more noise than previous works.
Expanded wavelength range and better fitting and model projection can provide improvements to the MODIS and VIIRS MTF assessments.

## Slide 20
Backup Slides
20

## Slide 21
Line Profile Correction Example
21
Surface variation past the edge depends strongly on the line profile location.
Edge curvature correction can be applied based on where the profile comes from.

## Slide 22
Solar Angle Dependence
22
The track direction MTF results show a dependence on solar angle (the scan direction less so).
Solar Angle = 76.6°
Solar Angle = 0.41°
Track Direction
Terra Band 1
SNPP Band I1
Seen in previous works as well.
Bad profiles near the terminator?
  Can use a filter if needed.
Edge curvature correction is applied but has minimal impact.