## Slide 1: JMA lunar calibration report [MTSAT-2 and GMS-5]
Masaya Takahashi1, Bartolomeo Viticchie2, and Sebastien Wagner2
1Japan Meteorological Agency, 2EUMETSAT
Acknowledgement: Tom Stone (USGS)

## Slide 2: Contents
Introduction
MTSAT-2/Imager lunar calibration
GMS-5/VISSR lunar calibration
Summary and future plans

## Slide 3: JMA geostationary (GEO) satellites
GMS (Geostationary Meteorological Satellite)
MTSAT (Multi-functional Transport SATellite )
Back-up operation of GMS-5 w/ GOES-9 by NOAA/NESDIS:
2003/05/22-2005/06/28
(GOES-9)
Aug 1981
GMS-2
(Himawari-2)
Aug 1984
GMS-3
(Himawari-3)
Sep 1989
GMS-4
(Himawari-4)
Mar 1995
GMS-5
(Himawari-5)
Feb 2005
MTSAT-1R
(Himawari-6)
Feb 2006
MTSAT-2
(Himawari-7)
Jul 1977
GMS
(Himawari)
Himawari
2014
Himawari-8
2016
Himawari-9
ⒸNASDA
ⒸNASDA
ⒸNASDA
ⒸNASDA
ⒸNASDA
ⒸSS/L
ⒸMELCO
Satellite | Operation period
GMS | 1978 – 1981
GMS-2 | 1981 – 1984
GMS-3 | 1984 – 1989
GMS-4 | 1989 – 1995
GMS-5 | 1995 – 2003
GOES-9 | 2003 – 2005
MTSAT-1R | 2005 – 2010
MTSAT-2 | 2010 – 2015
Himawari-8 | 2015 – 2022
Himawari-9 | 2022 – 2029
launched on
7 Oct. 2014!

## Slide 4: Lunar observations of JMA GEO satellites
Satellite | SSP Lon | H. Reso. of VIS/NIR at SSP | # of bands | Operation | Period used in this study | # of Obs. (for ROLO/total) | Phage angle
Himawari-8 | ~140 E | 0.5/1/2km | VIS:3 NIR:3 IR:10 | 2015 – 2022 (planned) | N/A at present |  | 
MTSAT-2 | 145 E | 1km | VIS:1 IR:4 | 2010 – 2015 | 2010 – 2013 | 62/98 | [-138,147]
MTSAT-1R | 140 E | 1km | VIS:1 IR:4 | 2005-2014 (Incl. back-up operation) | N/A due to its scanning pattern |  | 
GMS-5 | 140 E | 1.25km | VIS:1 IR:3 | 1995 – 2003 | 1995 – 2003 | 48/50 | [-94,96]
# of lunar observations for MTSAT-2 and GMS-5 will increase
Lunar observations for past GMS satellites will be extracted in future

## Slide 5: Spectral Response Functions
MTSAT-2/Imager
GMS-5/VISSR
MSG-3/SEVIRI

## Slide 6: MTSAT-2/Imager lunar calibration: general
3-axis stabilized GEO satellite
Satellite position: from telemetry
Located at ~ 145 E
Instrument: VIS/IR optical imager
Similar to GOES/Imager
1 visible channel (0.68µm)
1 km horizontal resolution at SSP
8 detectors for N-S direction
10-bit (1024 digital counts)
Scanning: E-W direction
Hourly Full-disk obs. + NH/SH obs.
The Moon acquisition
NO maneuver, unscheduled
2-5 lunar observations / month
2012-02-09T04:15Z
2012-10-24T22:00Z
2010-09-24T02:30Z
Visible channel calibration
Constant pre-launch values
Post launch cal. (not operationally used)
Vicarious calibration using RTM
Inter-calibration using deep convective clouds (DCCs)

## Slide 7
Manual extraction of lunar observations
Predict the Moon position using JPL DE405
IR (10.8µm) image is used to find the Moon
62 of total 98 observations for the ROLO model
(Jul. 2010-Dec.2013)
Lunar observation: L1A equivalent data
Converted to L1B equivalent DC/radiance (i.e., detector equalization using pre-launch values)
Oversampling: 1.75 in E-W direction
( = IFOV:28µrad / sampling width: 16µrad )
Dark signal correction: N/A
Computation of observables
Same method used in the GIRO imagette processing
Positions of the Moon wrt. the centre of the Earth for MTSAT-2/Imager (Jul.2010 – Dec.2013)
Earth disk
Absolute phase angle [deg]
Full moon
Low illumination
0 30 60 90 120 150 180
MTSAT-2/Imager lunar calibration: dataset

## Slide 8: ΔIrr (Calibration + Instrument drift)
Difference between the observed Irradiance and the ROLO Irradiance (in %)
Visible calibration coefficients for MTSAT-2/Imager
Constant pre-launch values -> ΔIrr shows only instrument drift
Time series of ΔIrr for MTSAT-2/Imager
Outliers
Δirr [%]
Hatch: 95% confidence interval of a linear regression line

## Slide 9: Outliers caused by incorrect extraction of lunar observations
Case #1
Case #2
Case #1: Careless extraction
Incomplete lunar disk image
Root cause of significant ΔIrr
VIS
IR1
(10.8 µm)
Case #2: Reasonable extraction at a glance
IR1
(10.8 µm)
VIS
East portion of the Moon: behind the Earth
Difficult to find out the Earth edge from VIS image during the night time

## Slide 10: Calibration slope vs. the ROLO (on the relative scale) - A way to monitor the instrument drift and compare with other calibration results
Δslope [%]
Lunar calibration
Good agreements with RTM and DCC, the smallest uncertainty of a regression line
Hatch: 95% confidence interval of a linear regression line
Time series of MTSAT-2/Imager calibration slope difference (each method is rescaled at the first calibration result)
Annual drift [%/yr]
DCC: 3.07 ± 0.16
Moon: 2.71 ± 0.15
RTM: 3.69 ± 0.62
FPIX : Oversampling factor
ΩPIX : Pixel solid angle [sr-1]
Dci : Digital count of the Moon pixel
OffDC : Averaged deep space DC
NPIX : Number of the Moon pixels
[W m-2 µm-1 DC-1]

## Slide 11: GMS-5/VISSR lunar calibration: general
Spin stabilized GEO satellite
Satellite position: from telemetry
Located at ~ 140 E
Instrument: VIS/IR optical imager
1 broad visible channel (similar to SEVIRI HRV)
1.25 km horizontal resolution at SSP
4 detectors for N-S direction
6-bit (64 digital counts)
Scanning: E-W direction
Hourly Full-disk observation
The Moon acquisition
NO maneuver, unscheduled
0-3 lunar observations / month
Visible channel calibration
Constant pre-launch values
Post launch cal. (not operationally used)
Vicarious calibration using RTM
Inter-calibration using deep convective clouds (DCCs)
Digital count Radiance
GMS-5/VISSR VIS 1996-05-02T00:53:40Z

## Slide 12
Absolute phase angle [deg]
Full moon
Low illumination
2 18 34 50 66 82 98
GMS-5/VISSR lunar calibration: dataset
Lunar observation: L1A equivalent data
Conversion of DC to radiance/reflectivity
Based on LUT, no calibration equation
Equalized DC does not exist
Radiance: derived using calibration coefficients for each detector
Oversampling: 1.475 for E-W direction
Dark signal correction: N/A
48 of total 50 useful lunar observations for the ROLO model (Jul. 1995-Feb. 2003)
Computation of observables
Radiance for the Moon masking
Positions of the Moon wrt. the centre of the Earth for GMS-5/VISSR
(Jul. 1995-Feb.2003)
Pixel number

## Slide 13: Digital counts handling in the GMS-5/VISSR lunar calibration
Sanity check file
Observed Irradiance: consistent, but not identical due to the Moon mask method
Radiance for “User”, DC for “Imgt Processing”
DC cannot be used because detectors are not equalized at count level
Observed DC, DC Offset and Total number of pixels: NO meaning in GMS-5/VISSR
Digital count Detector signal [V] Digital count
Reflectivity
Reflectivity
Detector signal [V]
(a) (b) (c)
GMS-5/VISSR VIS channel detector#2 signal-DC-reflectivity relationships
(Figures: Hashimoto et al. 2008)

## Slide 14: Phase angle dependence in ΔIrr
Difference between the observed Irradiance and the ROLO Irradiance (in %)
Clear phase angle dependence in ΔIrr may be caused by:
Incorrect computation of IrrOBS
Stray light effect, scan mirror angle dependence
ROLO model
…
Time series of ΔIrr for GMS-5/VISSR
Δirr [%]
Δirr [%]
The Moon phase angle [deg]

## Slide 15: Summary and future plans
Lunar calibration using the GIRO application
MTSAT-2/Imager
62 lunar observations (Jul. 2010 – Dec. 2013)
Calibration trend (annual drift) : consistent with other calibration method in terms of the uncertainty
GMS-5/VISSR
48 lunar observations (Jul. 1995 – Feb. 2003)
Need for the further investigation (phase angle dependence, detector equalization, usage of detector signals instead of digital counts, etc.)
Future plans
Application of GIRO apps. to Himawari-8/AHI 6 VIS/NIR bands
Combination of multiple calibration methods (DCC, desert, …)
Will be discussed/developed within the GSCIS framework

## Slide 16: Thank you for your attention

## Slide 17: Correction of phase angle dependence
Δirr [%]
The Moon phase angle [deg]
g : the Moon phase angle [deg]
ΔIrrCORR : ΔIrr w/ phase angle correction
GIRO (w/ phase angle correction)
Vicarious calibration using RTM and Terra/MODIS
ΔIrrCORR [%]
Drift: -0.58±0.11 [%/year]
Drift: -0.52±0.28 [%/year]
Good agreement in annual drift
Uncertainty is still large
w/o correction