## Slide 1
MODIS L1B Data Product Uncertainty Index Jack Xiong (Xiaoxiong.Xiong-1@nasa.gov), NASA GSFC
MODIS Characterization Support Team (MCST)
Instrument Background
Design Parameters and Primary Use of MODIS Spectral Bands
Moderate Resolution Imaging Spectroradiometer (MODIS)
On-board the Terra and Aqua spacecraft launched in December 1999 and May 2002, respectively.
36 spectral bands with a total of 490 individual detectors located on 4 focal plane assemblies (FPA).
  20 reflective solar bands (RSB): 1-19 and 26
  16 thermal emissive bands (TEB): 20-25 and 27-36
Three spatial resolutions (nadir): 250 m, 500 m, and 1 km
Scan angle: ±55° (from instrument nadir)
  A swath of 10 km (along-track) by 2330 km (nadir along-scan)
  Global coverage in less than 2 days
A broad range of applications
  Near 40 science data products for studies of the Earth’s land, ocean, and atmosphere properties.
Terra
Aqua
On-board calibrators
Solar diffuser (SD)
Solar diffuser stability monitor (SDSM)
Blackbody (BB)
Spectro-radiometric calibration assembly (SRCA)
Space view (SV) port
MODIS L1B Data Products
OBC |  | On-board Calibration and Telemetry Data
EV | QKM | 250 m Resolution RSB SI and UI (bands 1 and 2)
 | HKM | 500 m Resolution RSB SI and UI (bands 3-7) 250 m Aggregated 500 m RSB SI and UI (bands 1 and 2)
 | 1KM | 1 km Resolution RSB SI and UI (bands 8-19, 26) 1 km Resolution TEB SI and UI (bands 20-25, 27-36) 250 m Aggregated 1 km RSB SI and UI (bands 1 and 2) 500 m Aggregated 1 km RSB SI and UI (bands 3-7)
RSB reflectance factor
Reflectance = reflectance_scales * ( SI - reflectance_offsets )
RSB radiance
Radiance = Reflectance * Solar_Irradiance/( π *d2Earth-Sun )
TEB radiance
Radiance = radiance_scales * ( SI - radiance_offsets )
Uncertainty for both RSB and TEB
Uncertainty (%) = specified_uncertainty * exp( UI/scaling_factor )
reflectance_scales, reflectance_offsets, radiance_scales, radiance_offsets,
specified_uncertainty, and scaling_factor are band-specified attributes from each dataset at both native and aggregated resolutions.
Uncertainty and Uncertainty Index
Scale Integers (SI) | 
Value | Meaning
65524 - 65535 | Fill Values
65501 - 65523 | TBD
32768 - 65500 | NAD Closed Data
0-32767 | Normal Values
Uncertainty Index (UI) | 
Value | Meaning
0-14 | Normal UC Value
15 | UC > A Band Dependent Value; Non-calibratable SI
255 | Fill Values for Missing Data
NAD – Nadir Aperture Door UC - Uncertainty ;
MODIS L1B Detector QA Flags
Calibration Requirements:
±2% in reflectance and ±5% in radiance for RSB
±1% in radiance for most TEB, ±0.75% for band 20, ±10% for band 21, and ±0.5% for bands 31 and 32
Calibration requirements are specified at typical scene radiance levels and for observations within ±45º scan angle range
An additional 1% uncertainty is applied for the observations made at other scan angles and radiances from 0.3 specified typical radiance (0.3Ltyp) to 0.9 specified maximum radiance (0.9Lmax))
DQF Bit | Meaning
0 | Noisy Detector
1 | Dead Detector
2 | Out-of-Family Gain
3 | Dynamic Range
4 | Saturated DN
5 | High Fitting Residuals
6 | Crosstalk
7 | TBD
Detector Quality Flag (DQF)
Detector Quality Flag2 (DQF2)
DQF2 Bit | Meaning
0 | Noisy Sub-frame 1
1 | Noisy Sub-frame 2
2 | Noisy Sub-frame 3
3 | Noisy Sub-frame 4
4 | Dead Sub-frame 1
5 | Dead Sub-frame 2
6 | Dead Sub-frame 3
7 | Dead Sub-frame 4
490 detectors
180 detectors
Convert SI and UI to Calibrated Reflectance, Radiance, and Uncertainty
UI | Bands 1-4, 8-19 | Bands 5-7, 26 | Band 20 | Band 21 | Bands 22-25, 27-30, 33-36 | Bands 31, 32
0 | 1.50 | 1.50 | 0.56 | 2.50 | 0.50 | 0.38
1 | 1.73 | 1.83 | 0.69 | 3.21 | 0.64 | 0.48
2 | 2.00 | 2.24 | 0.84 | 4.12 | 0.82 | 0.62
3 | 2.30 | 2.73 | 1.02 | 5.29 | 1.06 | 0.79
4 | 2.66 | 3.34 | 1.25 | 6.80 | 1.36 | 1.02
5 | 3.06 | 4.08 | 1.53 | 8.73 | 1.75 | 1.31
6 | 3.53 | 4.98 | 1.87 | 11.20 | 2.24 | 1.68
7 | 4.08 | 6.08 | 2.28 | 14.39 | 2.88 | 2.16
8 | 4.70 | 7.43 | 2.79 | 18.47 | 3.69 | 2.77
9 | 5.43 | 9.07 | 3.40 | 23.72 | 4.74 | 3.56
10 | 6.26 | 11.08 | 4.16 | 30.46 | 6.09 | 4.57
11 | 7.22 | 13.54 | 5.08 | 39.11 | 7.82 | 5.87
12 | 8.33 | 16.53 | 6.20 | 50.21 | 10.04 | 7.53
13 | 9.61 | 20.20 | 7.57 | 64.48 | 12.90 | 9.67
14 | 11.08 | 24.67 | 9.25 | 82.79 | 16.56 | 12.42
15 | ≥ 12.79 | ≥ 30.13 | ≥ 11.30 | ≥106.30 | ≥ 21.26 | ≥ 15.95